## Position Log - Ball
Date | Time EDT | Operator Initials | Plateau | Final? | Target | RM1 |  | RM2 |  | RM3 |  | 5DOF X | 5DOF Y | 5DOF Z | 5DOF Rx | 5DOF Ry | AC Focus | Lidar Sig | Range (m) | Pressure (torr) | .fits | AC AZ (deg) | AC EL (deg) | X centr (pix) | Y centr (pix) | Notes Include ADM temp
|  |  | template | Y | sMATF mirror
|  |  | template | Y | sMATF LED pri
|  |  | template | Y | sMATF retro
|  |  | template | Y | CHIPS retro
|  |  | template | Y | Aperture mirror
|  |  | COPY & PASTE ONLY F COLUMN >> | Y | PATB mirror
|  |  | template | Y | PATB LED pri
|  |  | template | Y | PATB retro pri
|  |  | template | Y | PATA mirror
|  |  | template | Y | PATA LED pri
|  |  | template | Y | PATA retro pri
|  |  | closeout | Y | sMATF mirror
|  |  | closeout | Y | sMATF LED pri
|  |  | closeout | Y | sMATF retro
|  |  | closeout | Y | CHIPS retro
2023-05-17 | 6:00 | MW | Aperture wheel setup |  | Aperture mirror | 20.2274 | 23.8756 | 21.7523 | 23.5699 | 24.432 | 23.426 | [5 empty] | 0,28 |  |  |  | 92859 | 0.0083 | -0.18 |  |  | Initial alignment of aperture wheel mirror
2023-05-23 | 16:52 | MP | BA2 |  | sMATF mirror | 20.2274 | 23.8756 | 21.7523 | 23.5699 | 24.44 | 23.42 | [5 empty] | 0,18 | 14 | 4.0217 |  | 165228 | 0.0003 | 0.0049 |  |  | 1st pass of pre-shim ambient testing (5um stdev)
2023-05-23 | 16:57 | MP | BA2 |  | sMATF LED pri | 20.2274 | 23.8756 | 21.7523 | 23.5699 | 24.44 | 23.42 | [5 empty] | 1,75 |  |  |  | 165724 |  |  | 809.753 | 598.75 | Saved image has sMATF LED & bits together 165441
2023-05-23 | 17:05 | MP | BA2 |  | sMATF retro | 20.2274 | 23.8756 | 21.7523 | 23.5699 | 24.153 | 23.755 | [6 empty] | 15 | 4.0323 | [6 empty] | stdev 7.9 um
2023-05-23 | 17:12 | MP | BA2 |  | CHIPS retro | 20.2274 | 23.8756 | 21.7523 | 23.5699 | 24.63 | 23.18 | [6 empty] | 11 | 4.7633 | [6 empty] | stdev 4.2 um
2023-05-24 | 10:20 | MP | BA2 |  | sMATF mirror | 20.2274 | 23.8756 | 21.7523 | 23.5699 | 24.435 | 23.418 | [5 empty] | 0,25 | 11 | 4.0218 |  |  | -0.0053 | 0.0115 |  |  | 1st pass of pre-shim ambient testing (5um stdev)
2023-05-24 | 10:29 | MP | BA2 |  | sMATF LED pri | 20.2274 | 23.8756 | 21.7523 | 23.5699 | 24.44 | 23.42 | [5 empty] | 1,75 | [6 empty] | 813.61 | 598.59 | Saved image has sMATF LED & bits together 165441
2023-05-24 | 11:00 | MP | BA2 |  | sMATF retro | 20.2274 | 23.8756 | 21.7523 | 23.5699 | 24.44 | 23.42 | [6 empty] | 12 | 4.0323
2023-05-24 | 11:00 | MP | BA2 |  | CHIPS retro | 20.312 | 23.7254 | 21.5236 | 23.6494 | 24.63 | 23.18 | [6 empty] | 9 | 4.7633
2023-05-24 | 11:00 | MP | BA2 |  | sMATF mirror | 20.312 | 23.7254 | 21.5236 | 23.6494 | 24.44 | 23.42 | [5 empty] | 0,25 |  |  |  |  | 0.0017 | 0.0032
2023-05-24 | 11:34 | MP | BA2 |  | sMATF LED pri | 20.312 | 23.7254 | 21.5236 | 23.6494 | 24.44 | 23.42 | [5 empty] | 1,75 |  |  |  | 113229 |  |  | 807.523 | 603.562
2023-05-24 | 12:00 | MP | BA2 |  | sMATF mirror | 20.312 | 23.7254 | 21.5236 | 23.6494 | 24.44 | 23.42 | [5 empty] | 0,18 | 17 | 4.0217 |  | 115854 | -0.0022 | 0.0028
2023-05-24 | 11:44 | MP | BA2 |  | sMATF LED pri | 20.312 | 23.7254 | 21.5236 | 23.6494 | 24.44 | 23.42 | [5 empty] | 1,75 | [6 empty] | 807.48 | 601.85
2023-05-24 | 12:21 | MP | BA2 |  | sMATF mirror | 20.312 | 23.7254 | 21.5236 | 23.6494 | 24.44 | 23.42 | [5 empty] | 0,18 | 16 | 4.0217 |  | 122011 | -0.0091 | 0.0091
2023-05-24 | 12:31 | MP | BA2 |  | sMATF LED pri | 20.312 | 23.7254 | 21.5236 | 23.6494 | 24.44 | 23.42 | [5 empty] | 1,75 |  |  |  | 122920 |  |  | 813.655 | 595.84
2023-05-24 | 12:45 | MP | BA2 |  | sMATF mirror | 20.312 | 23.7254 | 21.5236 | 23.6494 | 24.44 | 23.42 | -2147 | -65 | -60.008 | 0 | -0.0005 | 0,18 |  |  |  | 124316 | -0.0094 | 0.0098
2023-05-24 | 12:42 | MP | BA2 |  | sMATF LED pri | 20.312 | 23.7254 | 21.5236 | 23.6494 | 24.44 | 23.42 | -2147 | -65 | -60.008 | 0 | -0.0005 | 1,72 |  |  |  | 124018 |  |  | 813.595 | 595.78
WOA-xxx pre shim
2023-05-24 | 13:27 | MP | BA2 | Y | sMATF mirror | 20.312 | 23.7254 | 21.5236 | 23.6494 | 24.44 | 23.42 | -2147 | -65 | -60.008 | 0 | -0.0005 | 0,18 | 13 | 4.0217 |  | 132329 | -0.0098 | 0.0101
2023-05-24 | 13:30 | MP | BA2 | Y | sMATF LED pri | 20.312 | 23.7254 | 21.5236 | 23.6494 | 24.44 | 23.42 | -2147 | -65 | -60.008 | 0 | -0.0005 | 1,73 |  |  |  | 132911 |  |  | 813.741 | 595.766
2023-05-24 | 13:34 | MP | BA2 | Y | sMATF retro | 20.312 | 23.7254 | 21.5236 | 23.6494 | 24.153 | 23.755 | -2147 | -65 | -60.008 | 0 | -0.0005 |  | 10 | 4.0322
2023-05-24 | 13:38 | MP | BA2 | Y | CHIPS retro | 20.312 | 23.7254 | 21.5236 | 23.6494 | 24.63 | 23.18 | -2147 | -65 | -60.008 | 0 | -0.0005 |  | 7 | 4.7633
2023-05-24 | 14:03 | MP | BA2 |  | PATB mirror | 20.312 | 23.7254 | 21.5236 | 23.6494 | 24.44 | 23.42 | 101.512 | 125.718 | 0 | 0.128 | 0.0822 | 0,18 |  |  |  |  | 0.0114 | 0.0188
2023-05-24 | 14:24 | MP | BA2 | Y | PATB mirror | 20.312 | 23.7254 | 21.5236 | 23.6494 | 24.44 | 23.42 | 101.512 | 125.718 | 0 | 0.1056 | 0.087038 | 0,18 | 18 | 3.4722 |  | 142345 | -0.0084 | 0.0098
2023-05-24 | 15:06 | MP | BA2 |  | PATB LED pri | 20.312 | 23.7254 | 21.5236 | 23.6494 | 24.44 | 23.42 | 88.615 | 108.614 | 0 | 0.1056 | 0.0871 | 1128 |  |  |  | 150627 |  |  | 805.636 | 589.993
2023-05-24 | 15:28 | MP | BA2 | Y | PATB LED pri | 20.312 | 23.7254 | 21.5236 | 23.6494 | 24.44 | 23.42 | 88.712 | 108.179 | 0 | 0.1056 | 0.0871 | 1128 |  |  |  | 152712 |  |  | 813.782 | 595.099
2023-05-24 | 16:09 | MP | BA2 | Y | PATB retro pri | 20.312 | 23.7254 | 21.5236 | 23.6494 | 24.44 | 23.42 | 105.901 | 99.041 | -0.001 | 0.1057 | 0.087 |  | 29 | 3.4586
2023-05-24 | 16:31 | MP | BA2 |  | PATA mirror | 20.312 | 23.7254 | 21.5236 | 23.6494 | 24.44 | 23.42 | -76.114 | 125.7 | -0.002 | 0.0768 | 0.0649 | 0,22 |  |  |  |  | -0.007 | -0.0048
2023-05-24 | 16:34 | MP | BA2 | Y | PATA mirror | 20.312 | 23.7254 | 21.5236 | 23.6494 | 24.44 | 23.42 | -76.114 | 125.7 | -0.002 | 0.083 | 0.0792 | 0,22 | 27 | 3.4722 |  | 163115 | -0.0098 | 0.0112
2023-05-24 | 16:59 | MP | BA2 | Y | PATA LED pri | 20.312 | 23.7254 | 21.5236 | 23.6494 | 24.44 | 23.42 | -60.334 | 108.121 | -0.003 | 0.0831 | 0.0793 | 1128 |  |  |  | 165732 |  |  | 812.892 | 594.974
2023-05-24 | 17:25 | MP | BA2 |  | PATA retro pri | 20.312 | 23.7254 | 21.5236 | 23.6494 | 24.44 | 23.42 | -72.08 | 98.978 | 0 | 0.0832 | 0.0791 |  | 11 | 3.4586
|  | MP | BA2 |  | Aperture mirror
|  | PM | BA2 (v2) |  | Aperture mirror | [20 empty] | 2nd pass of pre-shim ambient testing. Performed in opposite order as the 1st pass
2023-05-24 | 17:26 | MP | BA2 (v2) | Y | PATA retro pri | 20.312 | 23.7254 | 21.5236 | 23.6494 | 24.44 | 23.42 | -72.08 | 98.978 | 0 | 0.0832 | 0.0791 |  | 11 | 3.4586
2023-05-24 | 17:55 | MP | BA2 (v2) |  | PATA LED pri | 20.312 | 23.7254 | 21.5236 | 23.6494 | 24.44 | 23.42 | -60.333 | 108.12 | -0.002 | 0.0831 | 0.0791 | 1128 |  |  |  | 175412 |  |  | 811.924 | 595.721
2023-05-24 | 17:48 | MP | BA2 (v2) |  | PATA LED red | 20.312 | 23.7254 | 21.5236 | 23.6494 | 24.44 | 23.42 | -83.677 | 149.385 | 0.001 | 0.083 | 0.0787 | 1128 |  |  |  | 174619 |  |  | 795.779 | 581.584
2023-05-24 | 18:01 | MP | BA2 (v2) | Y | PATA LED pri | 20.312 | 23.7254 | 21.5236 | 23.6494 | 24.44 | 23.42 | -60.261 | 108.07 | -0.002 | 0.0831 | 0.0791 | 1128 |  |  |  | 175942 |  |  | 814.625 | 595.712
2023-05-24 | 18:08 | MP | BA2 (v2) | Y | PATA mirror | 20.312 | 23.7254 | 21.5236 | 23.6494 | 24.44 | 23.42 | -76.116 | 125.702 | -0.001 | 0.083 | 0.0789 | 0,19 | 28 | 3.4722 |  | 180546 | -0.0105 | 0.0119
2023-05-24 | 18:05 | MP | BA2 (v2) | Y | PATB retro pri | 20.312 | 23.7254 | 21.5236 | 23.6494 | 24.44 | 23.42 | 105.895 | 99.042 | 0 | 0.1054 | 0.0875 |  | 26 | 3.4586
2023-05-24 | 18:10 | MP | BA2 (v2) | Y | PATB LED pri | 20.312 | 23.7254 | 21.5236 | 23.6494 | 24.44 | 23.42 | 88.711 | 108.18 | 0 | 0.1056 | 0.0868 | 1128 |  |  |  | 182153 |  |  | 812.836 | 594.935 | The PAT LEDs appear to have drifted by ~1 pixel relative to the 1st pass
2023-05-24 | 18:25 | MP | BA2 (v2) | Y | PATB mirror | 20.312 | 23.7254 | 21.5236 | 23.6494 | 24.44 | 23.42 | 101.513 | 125.72 | 0 | 0.1055 | 0.0864 | 0,18 | 28 | 3.4722 |  | 183020 | -0.0098 | 0.0091
2023-05-24 | 18:52 | MP | BA2 (v2) | Y | CHIPS retro | 20.312 | 23.7254 | 21.5236 | 23.6494 | 24.63 | 23.18 | -2147.484 | -65 | -60.015 | -0.0004 | 0.002 |  | 6 | 4.7633
2023-05-24 | 18:56 | MP | BA2 (v2) | Y | sMATF retro | 20.312 | 23.7254 | 21.5236 | 23.6494 | 24.153 | 23.755 | -2147.484 | -65 | -60.015 | -0.0004 | 0.002 |  | 11 | 4.0323
2023-05-24 | 19:00 | MP | BA2 (v2) | Y | sMATF LED pri | 20.312 | 23.7254 | 21.5236 | 23.6494 | 24.44 | 23.42 | -2147.484 | -65 | -60.015 | -0.0004 | 0.002 | 1,75 |  |  |  | 185901 |  |  | 813.891 | 594.714 | Whereas the sMATF LED did not; implying the drift was from the 5DOF, not the ADM.
2023-05-24 | 19:02 | MP | BA2 (v2) | Y | sMATF mirror | 20.312 | 23.7254 | 21.5236 | 23.6494 | 24.44 | 23.42 | -2147.484 | -65 | -60.015 | -0.0004 | 0.002 | 0,19 | 13 | 4.0217 |  | 190116 | -0.0098 | 0.0108 |  |  | << Last RM positions prior to shim
2023-05-25 | 9:11 | MP | BA3 |  | sMATF mirror | 20.312 | 23.7254 | 21.5236 | 23.6494 | 24.44 | 23.41 | -2147.484 | -65 | -60.015 | -0.0004 | 0.002 | 0,24 | 4 | 4.022 |  | 90955 | -0.0063 | 0.0115 |  |  | Initial meas after shim
2023-05-25 | 9:22 | MP | BA3 |  | sMATF LED pri | 20.312 | 23.7254 | 21.5236 | 23.6494 | 24.44 | 23.41 | -2147.484 | -65.002 | -60.01 | -0.0001 | 0.0009 | 1,77 |  |  |  | 92211 |  |  | 811.799 | 588.529 | Initial meas after shim
2023-05-25 | 10:24 | MP | BA3 |  | sMATF mirror | 20.312 | 23.7254 | 21.5236 | 23.6494 | 24.44 | 23.41 | -2147.484 | -65.002 | -60.01 | -0.0001 | 0.0009 | 0,20 |  |  |  | 102436 | 0.0107 | 0.0049 |  |  | Nulling
2023-05-25 | 10:15 | MP | BA3 |  | sMATF LED pri | 20.2428 | 23.757 | 21.568 | 23.593 | 24.432 | 23.418 | -2147.484 | -65.002 | -60.01 | -0.0001 | 0.0009 | 1,77 | [6 empty] | 807.598 | 601.839 | Nulling
2023-05-25 | 10:39 | MP | BA3 |  | sMATF mirror | 20.2428 | 23.757 | 21.568 | 23.593 | 24.432 | 23.418 | -2147.484 | -65.002 | -60.01 | -0.0001 | 0.0009 | 0,23 | 18 | 4.022 |  | 103750 | 0.0031 | 0.0039 |  |  | Post null
2023-05-25 | 10:33 | MP | BA3 |  | sMATF LED pri | 20.24 | 23.7565 | 21.5675 | 23.5927 | 24.432 | 23.418 | -2147.484 | -65.002 | -60.01 | -0.0001 | 0.0009 | 1,74 |  |  |  | 103325 |  |  | 807.86 | 601.839 | Post null
2023-05-25 | 11:01 | MP | BA3 |  | sMATF mirror | 20.2428 | 23.757 | 21.568 | 23.593 | 24.432 | 23.418 | -2147.484 | -65.002 | -60.01 | -0.0001 | 0.0009 | 0,23 | 10 | 4.0219 |  | 110054 | 0.0006 | 0.0042 |  |  | (monitoring drift)
2023-05-25 | 11:24 | MP | BA3 |  | sMATF LED pri | 20.24 | 23.7565 | 21.5675 | 23.5927 | 24.432 | 23.418 | -2147.484 | -65.002 | -60.01 | -0.0001 | 0.0009 | 1,74 |  |  |  | 112148 |  |  | 809.017 | 599.68 | disabled motors at 11:28
2023-05-25 | 11:52 | MP | BA3 |  | sMATF mirror | 20.2428 | 23.757 | 21.568 | 23.593 | 24.432 | 23.418 | -2147.484 | -65.002 | -60.01 | -0.0001 | 0.0009 | 0,19 | 16 | 4.022 |  | 115202 | -0.0001 | 0.007 |  |  | (monitoring drift)
2023-05-25 | 11:49 | MP | BA3 |  | sMATF LED pri | 20.24 | 23.7565 | 21.5675 | 23.5927 | 24.432 | 23.418 | -2147.484 | -65.002 | -60.01 | -0.0001 | 0.0009 | 1,74 |  |  |  | 114910 |  |  | 811.8 | 598.7
2023-05-25 | 12:12 | MP | BA3 |  | sMATF mirror | 20.2428 | 23.757 | 21.568 | 23.593 | 24.432 | 23.418 | -2147.484 | -65.002 | -60.01 | -0.0001 | 0.0009 | 0,19 |  |  |  | 121252 | -0.0001 | 0.0077 |  |  | (monitoring drift)
2023-05-25 | 12:16 | MP | BA3 |  | sMATF LED pri | 20.24 | 23.7565 | 21.5675 | 23.5927 | 24.432 | 23.418 | -2147.484 | -65.002 | -60.01 | -0.0001 | 0.0009 | 1,74 |  |  |  | 121605 |  |  | 811.9 | 597.8
2023-05-25 | 12:34 | MP | BA3 |  | sMATF mirror | 20.2428 | 23.757 | 21.568 | 23.593 | 24.432 | 23.418 | -2147.484 | -65.002 | -60.01 | -0.0001 | 0.0009 | 0,21 | 15 | 4.022 |  | 123337 | 0.0006 | 0.0081 |  |  | (monitoring drift)
2023-05-25 | 12:31 | MP | BA3 |  | sMATF LED pri | 20.24 | 23.7565 | 21.5675 | 23.5927 | 24.432 | 23.418 | -2147.484 | -65.002 | -60.01 | -0.0001 | 0.0009 | 1,74 |  |  |  | 123144 |  |  | 812.6 | 598.6
2023-05-25 | 12:49 | MP | BA3 |  | sMATF mirror | 20.2428 | 23.757 | 21.568 | 23.593 | 24.432 | 23.418 | -2147.484 | -65.002 | -60.01 | -0.0001 | 0.0009 | 0,21 | 14 | 4.022 |  | 124959 | 0.0006 | 0.0084 |  |  | (monitoring drift)
2023-05-25 | 12:52 | MP | BA3 |  | sMATF LED pri | 20.24 | 23.7565 | 21.5675 | 23.5927 | 24.432 | 23.418 | -2147.484 | -65.002 | -60.01 | -0.0001 | 0.0009 | 1,74 |  |  |  | 125216 |  |  | 812.6 | 598.7 | 0.8 pix drift/1hr
2023-05-25 | 13:02 | MP | BA3 | Y | sMATF retro | 20.24 | 23.7565 | 21.5675 | 23.5927 | 24.145 | 23.755 | -2147.484 | -65.002 | -60.01 | -0.0001 | 0.0009 |  | 11 | 4.0325
2023-05-25 | 13:07 | MP | BA3 |  | CHIPS retro | 20.24 | 23.7565 | 21.5675 | 23.5927 | 24.63 | 23.18 | -2147.484 | -65.002 | -60.01 | -0.0001 | 0.0009 |  | 7 | 4.7636 | [6 empty] | 13:15 PURGE LINE INSTALL
2023-05-25 | 13:52 | MP | BA3 | Y | sMATF mirror | 20.24 | 23.7565 | 21.5675 | 23.5927 | 24.432 | 23.418 | -2147.484 | -65.002 | -60.01 | -0.0001 | 0.0009 | 0,20 | 13 | 4.022 |  | 135246 | 0.0006 | 0.0087 |  |  | After purge line install
2023-05-25 | 13:56 | MP | BA3 | Y | sMATF LED pri | 20.24 | 23.7565 | 21.5675 | 23.5927 | 24.432 | 23.418 | -2147.484 | -65.002 | -60.01 | -0.0001 | 0.0009 | 1,74 |  |  |  | 135632 |  |  | 812.8 | 598.05
2023-05-25 | 13:56 | MP | BA3 | Y | CHIPS retro | 20.24 | 23.7565 | 21.5675 | 23.5927 | 24.63 | 23.18 | -2147.484 | -65.002 | -60.01 | -0.0001 | 0.0009 |  | 8 | 4.7635
2023-05-25 | 14:29 | MP | BA3 | Y | PATB mirror | 20.24 | 23.7565 | 21.5675 | 23.5927 | 24.432 | 23.42 | 101.513 | 125.721 | -0.002 | 0.1052 | 0.0864 | 0,21 | 23 | 3.4722 |  | 142851 | -0.0001 | 0.0063 |  |  | Used jog buttons to move laser spot in camera. Go to the old positions did not seem to repeat alignment. Moved laser spot to encoder positions shown here.
2023-05-25 | 14:50 | MP | BA3 |  | PATB LED pri | 20.24 | 23.7565 | 21.5675 | 23.5927 | 24.432 | 23.42 | 88.712 | 108.18 | -0.002 | 0.1056 | 0.086 | 1133 | [6 empty] | 813.1 | 609.6
2023-05-25 | 15:02 | MP | BA3 | Y | PATB LED pri | 20.24 | 23.7565 | 21.5675 | 23.5927 | 24.432 | 23.42 | 88.977 | 108.621 | -0.002 | 0.1056 | 0.086 | 1133 |  |  |  | 150013 |  |  | 812.7 | 598.8
2023-05-25 | 15:08 | MP | BA3 | Y | PATB retro pri | 20.24 | 23.7565 | 21.5675 | 23.5927 | 24.432 | 23.42 | 105.9 | 99.042 | 0 | 0.1052 | 0.0874 |  | 21 | 3.4586
2023-05-25 | 15:14 | MP | BA3 | Y | PATA mirror | 20.24 | 23.7565 | 21.5675 | 23.5927 | 24.432 | 23.42 | -76.116 | 125.702 | -0.001 | 0.083 | 0.0789 | 0,21 | 22 | 3.4722 |  | 151416 | -0.0008 | 0.0084
2023-05-25 | 15:20 | MP | BA3 |  | PATA LED pri | 20.24 | 23.7565 | 21.5675 | 23.5927 | 24.432 | 23.42 | -60.261 | 108.07 | -0.002 | 0.0831 | 0.0791 | 1136 |  |  |  | 152046 |  |  | 815.7 | 608.6
2023-05-25 | 15:25 | MP | BA3 | Y | PATA LED pri | 20.24 | 23.7565 | 21.5675 | 23.5927 | 24.432 | 23.42 | -60.113 | 108.543 | -0.002 | 0.0831 | 0.0791 | 1136 |  |  |  | 152332 |  |  | 812.7 | 598.7
2023-05-25 | 15:28 | MP | BA3 | Y | PATA retro pri | 20.24 | 23.7565 | 21.5675 | 23.5927 | 24.432 | 23.42 | -72.082 | 98.978 | 0.001 | 0.083 | 0.0796 |  | 11 | 3.4586
2023-05-26 | 10:45 | MP | 5/26 BA4 |  | sMATF mirror | 20.2408 | 23.7535 | 21.5675 | 23.5927 | 24.432 | 23.42 | -2147.484 | -65 | -60.007 | 0 | 0 | 0,20 |  |  |  | 104509 | -0.0654 | -0.06 |  |  | Initial meas after shim
2023-05-26 | 10:45 | MP | 5/26 BA4 |  | sMATF LED pri | 20.2408 | 23.7535 | 21.5675 | 23.5927 | 24.432 | 23.42 | -2147.484 | -65 | -60.007 | 0 | 0 | 1,75 |  |  |  | 104305 |  |  | 807.9 | 602.6 | Initial meas after shim
2023-05-26 | 10:58 | MP | 5/26 BA4 |  | sMATF mirror | 20.6158 | 23.3765 | 21.0305 | 23.9393 | 24.432 | 23.42 | -2147.484 | -65 | -60.007 | 0 | 0 | 0,21 |  |  |  |  | -0.0334 | -0.0263 |  |  | adj 1
2023-05-26 | 10:55 | MP | 5/26 BA4 |  | sMATF LED pri | 20.6158 | 23.3765 | 21.0305 | 23.9393 | 24.432 | 23.42 | -2147.484 | -65 | -60.007 | 0 | 0 | 1,76 |  |  |  | 105515 |  |  | 807.8 | 601.7 | adj 1
2023-05-26 | 11:16 | MP | 5/26 BA4 |  | sMATF mirror | 20.8118 | 23.2155 | 20.7951 | 24.1163 | 24.432 | 23.42 | -2147.484 | -65 | -60.007 | 0 | 0 | 0,21 |  |  |  | 111622 | -0.0188 | -0.0097 |  |  | adj 2
2023-05-26 | 11:12 | MP | 5/26 BA4 |  | sMATF LED pri | 20.8118 | 23.2155 | 20.7951 | 24.1163 | 24.432 | 23.42 | -2147.484 | -65 | -60.007 | 0 | 0 | 1,76 |  |  |  | 111239 |  |  | 807.7 | 601.74
2023-05-26 | 11:37 | MP | 5/26 BA4 |  | sMATF mirror | 20.9817 | 23.1325 | 20.6689 | 24.2679 | 24.432 | 23.42 | -2147.484 | -65 | -60.007 | 0 | 0 | 0,21 |  |  |  | 113729 | -0.0049 | -0.0013 |  |  | Nulled here, starting settle clock
2023-05-26 | 11:35 | MP | 5/26 BA4 |  | sMATF LED pri | 20.9817 | 23.1325 | 20.6689 | 24.2679 | 24.432 | 23.42 | -2147.484 | -65 | -60.007 | 0 | 0 | 1,76 |  |  |  | 113540 |  |  | 807.7 | 601.9
2023-05-26 | 11:53 | MP | 5/26 BA4 |  | sMATF mirror | 20.9817 | 23.1325 | 20.6689 | 24.2679 | 24.432 | 23.42 | -2147.484 | -65 | -60.007 | 0 | 0 | 0,21 |  |  |  | 115331 | -0.0042 | -0.0027 |  |  | settle check
2023-05-26 | 11:55 | MP | 5/26 BA4 |  | sMATF LED pri | 20.9817 | 23.1325 | 20.6689 | 24.2679 | 24.432 | 23.42 | -2147.484 | -65 | -60.007 | 0 | 0 | 1,76 |  |  |  | 115506 |  |  | 806.1 | 601.8 | settle check
2023-05-26 | 12:17 | MP | 5/26 BA4 |  | sMATF mirror | 20.9817 | 23.1325 | 20.6689 | 24.2679 | 24.432 | 23.42 | -2147.484 | -65 | -60.007 | 0 | 0 | 0,21 |  |  |  | 121708 | -0.0042 | -0.0027 |  |  | settle check
2023-05-26 | 12:15 | MP | 5/26 BA4 |  | sMATF LED pri | 20.9817 | 23.1325 | 20.6689 | 24.2679 | 24.432 | 23.42 | -2147.484 | -65 | -60.007 | 0 | 0 | 1,76 |  |  |  | 121519 |  |  | 806.75 | 601.745 | settle check
2023-05-26 | 12:43 | MP | 5/26 BA4 |  | sMATF mirror | 20.9817 | 23.1325 | 20.6689 | 24.2679 | 24.432 | 23.42 | -2147.484 | -65 | -60.007 | 0 | 0 | 0,21 |  |  |  | 124316 | -0.0039 | -0.0027 |  |  | settle check
2023-05-26 | 12:46 | MP | 5/26 BA4 |  | sMATF LED pri | 20.9817 | 23.1325 | 20.6689 | 24.2679 | 24.432 | 23.42 | -2147.484 | -65 | -60.007 | 0 | 0 | 1,77 |  |  |  | 124610 |  |  | 806.8 | 602.5 | settle check
2023-05-26 | 12:59 | MP | 5/26 BA4 | Y | sMATF mirror | 20.9817 | 23.1325 | 20.6689 | 24.2679 | 24.422 | 23.42 | -2147.484 | -65 | -60.007 | 0 | 0 | 0,20 | 6 | 4.0215 |  | 125935 | -0.0035 | -0.0027
2023-05-26 | 12:57 | MP | 5/26 BA4 | Y | sMATF LED pri | 20.9817 | 23.1325 | 20.6689 | 24.2679 | 24.432 | 23.41 | -2147.484 | -65 | -60.007 | 0 | 0 | 1,77 |  |  |  | 125746 |  |  | 806 | 601.8
2023-05-26 | 13:14 | MP | 5/26 BA4 | Y | sMATF retro | 20.9817 | 23.1325 | 20.6689 | 24.2679 | 24.115 | 23.72 | -2147.484 | -65 | -60.007 | 0 | 0 |  | 28 | 4.032
2023-05-26 | 13:16 | MP | 5/26 BA4 | Y | CHIPS retro | 20.9817 | 23.1325 | 20.6689 | 24.2679 | 24.63 | 23.18 | -2147.484 | -65 | -60.007 | 0 | 0 |  | 13 | 4.7631
2023-05-26 | 13:38 | MP | 5/26 BA4 |  | PATB mirror | 20.9817 | 23.1325 | 20.6689 | 24.2679 | 24.422 | 23.42 | 101.5 | 125.722 | -0.004 | 0.1053 | 0.0864 | 0,20 |  |  |  | 133805 | 0.0576 | 0.0674
2023-05-26 | 14:22 | MP | 5/26 BA4 | Y | PATB mirror | 20.9817 | 23.1325 | 20.6689 | 24.2679 | 24.435 | 23.42 | 101.5 | 125.722 | 0 | 0.0172 | 0.0661 | 0,20 | 22 | 3.4722 |  | 142220 | -0.0042 | -0.0048
2023-05-26 | 14:38 | MP | 5/26 BA4 |  | PATB LED pri | 20.9817 | 23.1325 | 20.6689 | 24.2679 | 24.435 | 23.42 | 88.977 | 108.622 | 0 | 0.0172 | 0.0658 | 1133 | [6 empty] | 803.9 | 580.8
2023-05-26 | 14:45 | MP | 5/26 BA4 | Y | PATB LED pri | 20.9817 | 23.1325 | 20.6689 | 24.2679 | 24.435 | 23.42 | 88.49 | 107.737 | 0 | 0.0172 | 0.0658 | 1133 |  |  |  | 144550 |  |  | 805.7 | 601.7
2023-05-26 | 14:53 | MP | 5/26 BA4 | Y | PATB retro pri | 20.9817 | 23.1325 | 20.6689 | 24.2679 | 24.435 | 23.42 | 105.9 | 99.042 | 0 | 0.0172 | 0.0658 |  | 20 | 3.4586
2023-05-26 | 15:04 | MP | 5/26 BA4 |  | PATA mirror | 20.9817 | 23.1325 | 20.6689 | 24.2679 | 24.435 | 23.42 | -76.116 | 125.702 | -0.001 | 0.083 | 0.0789 | 0,20 |  |  |  | 150410 | 0.0565 | 0.0674 |  |  | GSA move did not converge on PATB mirror 18:58:25 UTC
2023-05-26 | 15:05 | MP | 5/26 BA4 | Y | PATA mirror | 20.9817 | 23.1325 | 20.6689 | 24.2679 | 24.435 | 23.42 | -76.116 | 125.702 | -0.001 | -0.0061 | 0.055 | 0,20 | 15 | 3.4723 |  | 150530 | -0.0035 | -0.001
2023-05-26 | 15:11 | MP | 5/26 BA4 |  | PATA LED pri | 20.9817 | 23.1325 | 20.6689 | 24.2679 | 24.435 | 23.42 | -60.261 | 108.07 | -0.002 | -0.0061 | 0.055 | 1133 |  |  |  | 151146 |  |  | 803.8 | 580.7
2023-05-26 | 15:14 | MP | 5/26 BA4 | Y | PATA LED pri | 20.9817 | 23.1325 | 20.6689 | 24.2679 | 24.435 | 23.42 | -60.584 | 107.65 | -0.002 | -0.0061 | 0.055 | 1133 |  |  |  | 151446 |  |  | 805.7 | 601.7
2023-05-26 | 15:20 | MP | 5/26 BA4 | Y | PATA retro pri | 20.9817 | 23.1325 | 20.6689 | 24.2679 | 24.435 | 23.42 | -72.082 | 98.978 | 0.001 | 0.083 | 0.0796 |  | 10 | 3.4586
2023-05-26 | 15:39 | MP | 5/26 BA4 v2 | Y | sMATF mirror | 20.9817 | 23.1325 | 20.6689 | 24.2679 | 24.435 | 23.41 | -2147.484 | -65 | -60.004 | 0 | 0 | 0,20 | 11 | 4.0215 |  | 153929 | -0.0122 | 0.0056
2023-05-26 | 15:48 | MP | 5/26 BA4 v2 | Y | sMATF LED pri | 20.9817 | 23.1325 | 20.6689 | 24.2679 | 24.435 | 23.42 | -2147.484 | -65 | -60.004 | 0 | 0 | 1,77 |  |  |  | 161519 | -0.0087 | 0.0083 | 813.7 | 595.6 | The sMATF LED appears to have shifted! But the RMs were powered off and nobody was in the cleanroom. The prior run is very suspect
2023-05-26 | 16:20 | MP | 5/26 BA4 v2 | Y | sMATF retro | 20.9817 | 23.1325 | 20.6689 | 24.2679 | 24.12 | 23.72 | -2147.484 | -65 | -60.004 | 0 | 0 |  | 15 | 4.032 |  |  |  |  | 7.7 | -6.2
2023-05-26 | 16:35 | MP | 5/26 BA4 v2 |  | PATB mirror | 20.9817 | 23.1325 | 20.6689 | 24.2679 | 24.435 | 23.42 | 101.513 | 125.725 | 0 | 0.0167 | 0.0656 | 0,20 | 30 | 3.4722 |  | 163500 | -0.0053 | -0.0013 |  |  | GSM move did not converge 20:35:16 UTC
2023-05-26 | 16:49 | MP | 5/26 BA4 v2 |  | PATB LED pri | 20.9817 | 23.1325 | 20.6689 | 24.2679 | 24.435 | 23.42 | 88.489 | 107.737 | 0 | 0.0167 | 0.0656 | 1128 |  |  |  | 164951 |  |  | 805.95 | 600.93
2023-05-26 | 17:03 | MP | 5/26 BA4 v2 |  | PATB mirror | 20.9817 | 23.1325 | 20.6689 | 24.2679 | 24.435 | 23.42 | 101.513 | 125.725 | 0 | 0.0173 | 0.0658 | 0,20 |  |  |  | 170317 | -0.0049 | 0.0008
2023-05-26 | 17:04 | MP | 5/26 BA4 v2 | Y | PATB mirror | 20.9817 | 23.1325 | 20.6689 | 24.2679 | 24.435 | 23.42 | 101.513 | 125.725 | 0 | 0.0144 | 0.0738 | 0,20 | 32 | 3.4723 |  | 170444 | -0.0105 | 0.0049
2023-05-26 | 17:16 | MP | 5/26 BA4 v2 | Y | PATB LED pri | 20.9817 | 23.1325 | 20.6689 | 24.2679 | 24.435 | 23.42 | 88.936 | 107.73 | 0 | 0.0143 | 0.0735 | 1130 |  |  |  | 171629 |  |  | 813.78 | 596.6
2023-05-26 | 17:25 | MP | 5/26 BA4 v2 |  | PATB retro pri | 20.9817 | 23.1325 | 20.6689 | 24.2679 | 24.435 | 23.42 | 105.9 | 99.042 | 0 | 0.0145 | 0.0741 |  | 23 | 3.4586
2023-05-26 | 17:37 | MP | 5/26 BA4 v2 |  | sMATF mirror | 20.9817 | 23.1325 | 20.6689 | 24.2679 | 24.435 | 23.41 | -2147.484 | -65 | -60.004 | 0 | 0 | 0,20 | 14 | 4.0215 |  | 173715 | -0.0126 | 0.0063
2023-05-26 | 17:44 | MP | 5/26 BA4 v2 |  | sMATF LED pri | 20.9817 | 23.1325 | 20.6689 | 24.2679 | 24.435 | 23.42 | -2147.484 | -65 | -60.004 | 0 | 0 | 1,76 |  |  |  | 174422 |  |  | 814.7 | 594.6
2023-05-26 | 17:50 | MP | 5/26 BA4 v2 |  | sMATF retro | 20.9817 | 23.1325 | 20.6689 | 24.2679 | 24.12 | 23.72 | -2147.484 | -65 | -60.004 | 0 | 0 |  | 15 | 4.032
2023-05-30 | 11:47 | MP | 5/30 BA4 | Y | sMATF mirror | 20.9817 | 23.1325 | 20.6689 | 24.2679 | 24.435 | 23.41 | -2147.484 | -65 | -60.004 | 0 | 0 | 0,20 | 27 | 4.0215 |  | 114745 | -0.0126 | 0.007 |  |  | Do not re-null RMs
2023-05-30 | 11:50 | MP | 5/30 BA4 | Y | sMATF LED pri | 20.9817 | 23.1325 | 20.6689 | 24.2679 | 24.435 | 23.41 | -2147.484 | -65 | -60.004 | 0 | 0 | 1,76 |  | -0.010549 |  | 115044 | 0 | 0.0007 | 814.59 | 594.7
2023-05-30 | 12:01 | MP | 5/30 BA4 | Y | sMATF retro | 20.9817 | 23.1325 | 20.6689 | 24.2679 | 24.12 | 23.72 | -2147.484 | -65 | -60.004 | 0 | 0 |  | 11 | 4.032 |  |  |  |  | -0.11 | 0.1 | (Retro ranged prior to mirror measurement since RM3 was still at Friday's RM3 position)
2023-05-30 | 12:06 | MP | 5/30 BA4 | Y | CHIPS retro | 20.9817 | 23.1325 | 20.6689 | 24.2679 | 24.63 | 23.18 | -2147.484 | -65 | -60.004 | 0 | 0 |  | 10 | 4.7631
2023-05-30 | 12:29 | MP | 5/30 BA4 | Y | PATB mirror | 20.9817 | 23.1325 | 20.6689 | 24.2679 | 24.435 | 23.41 | 101.513 | 125.725 | 0 | 0.0143 | 0.0737 | 0,20 | 13 | 3.4722 |  | 122925 | -0.0119 | 0.0056
2023-05-30 | 12:46 | MP | 5/30 BA4 | Y | PATB LED pri | 20.9817 | 23.1325 | 20.6689 | 24.2679 | 24.435 | 23.41 | 88.936 | 107.73 | 0 | 0.0143 | 0.0735 | 1130 |  | -5e-05 |  | 124649 | 0.0007 | -0.0007 | 813.9 | 595.8
2023-05-30 | 13:07 | MP | 5/30 BA4 |  | sMATF mirror | 20.9817 | 23.1325 | 20.6689 | 24.2679 | 24.437 | 23.414 | -2147.484 | -65 | -60.004 | 0 | 0 | 0,20 | 11 | 4.0215 |  | 130739 | -0.0133 | 0.0063 | 0.12 | -0.8 | << differences wrt Friday
2023-05-30 | 13:16 | MP | 5/30 BA4 |  | sMATF LED pri | 20.9817 | 23.1325 | 20.6689 | 24.2679 | 24.437 | 23.414 | -2147.484 | -65 | -60.004 | 0 | 0 | 1,76 |  |  |  | 131613 | -0.0007 | -0.0007 | 814.7 | 593.8
2023-05-30 | 13:35 | MP | 5/30 BA4 |  | PATB mirror | 20.9817 | 23.1325 | 20.6689 | 24.2679 | 24.437 | 23.414 | 101.513 | 125.725 | 0 | 0.0143 | 0.0737 | 0,20 | 18 | 3.4722 |  | 133500 | -0.0115 | 0.0056 | 0 | -0.8
2023-05-30 | 13:44 | MP | 5/30 BA4 |  | PATB LED pri | 20.9817 | 23.1325 | 20.6689 | 24.2679 | 24.437 | 23.414 | 88.936 | 107.73 | 0 | 0.0143 | 0.0735 | 1130 |  |  |  | 134406 | 0.0004 | 0 | 813.95 | 594.8
2023-05-30 | 14:00 | MP | 5/30 BA4 |  | sMATF mirror | 20.9817 | 23.1325 | 20.6689 | 24.2679 | 24.437 | 23.414 | -2147.484 | -65 | -60.004 | 0 | 0 | 0,20 | 11 | 4.0215 |  | 140020 | -0.0133 | 0.0063 | 0.17 | -1.8
2023-05-30 | 14:03 | MP | 5/30 BA4 |  | sMATF LED pri | 20.9817 | 23.1325 | 20.6689 | 24.2679 | 24.437 | 23.414 | -2147.484 | -65 | -60.004 | 0 | 0 | 1,76 |  |  |  | 140317 | 0 | 0 | 814.75 | 593.6
2023-05-30 | 14:19 | MP | 5/30 BA4 |  | PATB mirror | 20.9817 | 23.1325 | 20.6689 | 24.2679 | 24.437 | 23.414 | 101.513 | 125.725 | 0 | 0.0143 | 0.0737 | 0,20 | 17 | 3.4722 |  | 141918 | -0.0119 | 0.0056 | 0.05 | -1
2023-05-30 | 14:29 | MP | 5/30 BA4 |  | PATB LED pri | 20.9817 | 23.1325 | 20.6689 | 24.2679 | 24.437 | 23.414 | 88.936 | 107.73 | 0 | 0.0143 | 0.0735 | 1130 |  |  |  | 142922 |  |  | 814 | 594.98 | (14:45-15:30): camera was ON thru lunch
2023-05-30 | 16:09 | MP | 5/30 BA4 |  | sMATF mirror | 20.9817 | 23.1325 | 20.6689 | 24.2679 | 24.437 | 23.414 | -2147.484 | -65 | -60.004 | 0 | 0 | 0,20 | 15 | 4.0215 |  | 160922 | -0.014 | 0.006 | 0.22 | -1.62 | 16:05 Pat & Doug went in to remove piece of paper left in beam
2023-05-30 | 16:12 | MP | 5/30 BA4 |  | sMATF LED pri | 20.9817 | 23.1325 | 20.6689 | 24.2679 | 24.437 | 23.414 | -2147.484 | -65 | -60.004 | 0 | 0 | 1,76 |  |  |  | 161225 |  |  | 814.7 | 593.6
2023-05-30 | 16:30 | MP | 5/30 BA4 |  | PATB mirror | 20.9817 | 23.1325 | 20.6689 | 24.2679 | 24.437 | 23.414 | 101.513 | 125.725 | 0 | 0.0143 | 0.0737 | 0,20 | 22 | 3.4722 |  | 163037 | -0.0133 | 0.0056 | 0 | -1
2023-05-30 | 16:43 | MP | 5/30 BA4 |  | PATB LED pri | 20.9817 | 23.1325 | 20.6689 | 24.2679 | 24.437 | 23.414 | 88.936 | 107.73 | 0 | 0.0143 | 0.0735 | 1130 |  |  |  | 164312 |  |  | 813.9 | 594.9
2023-05-30 | 16:59 | MP | 5/30 BA4 |  | sMATF mirror | 20.9817 | 23.1325 | 20.6689 | 24.2679 | 24.437 | 23.414 | -2147.484 | -65 | -60.004 | 0 | 0 | 0,20 | 13 | 4.0215 |  | 165927 | -0.0136 | 0.006 | 0.12 | -1.7 | #5
2023-05-30 | 17:03 | MP | 5/30 BA4 |  | sMATF LED pri | 20.9817 | 23.1325 | 20.6689 | 24.2679 | 24.437 | 23.414 | -2147.484 | -65 | -60.004 | 0 | 0 | 1,76 |  |  |  | 170328 |  |  | 814.9 | 593.1 | ** had to turn sMATF LED down to 2mA
2023-05-30 | 17:08 | MP | 5/30 BA4 |  | sMATF retro | 20.9817 | 23.1325 | 20.6689 | 24.2679 | 24.12 | 23.72 | -2147.484 | -65 | -60.004 | 0 | 0 |  | 19 | 4.032 |  |  |  |  | 0.2 | -1.5
2023-05-30 | 17:10 | MP | 5/30 BA4 |  | CHIPS retro | 20.9817 | 23.1325 | 20.6689 | 24.2679 | 24.63 | 23.18 | -2147.484 | -65 | -60.004 | 0 | 0 |  | 11 | 4.7631
2023-05-31 | 12:10 | MP | 5/31 nulling |  |  | 20.312 | 23.7254 | 21.5236 | 23.6494 | 24.44 | 23.42 | [14 empty] | RMs set to last pre-shim position (5/24 19:02 EDT) RMs reset to yesterdays position (5/30 16:59 EDT)
2023-05-31 | 20:53 | MP | 5/31 nulling |  | sMATF mirror | 20.9858 | 23.1351 | 20.6689 | 24.2679 | 24.437 | 23.414 | -2147.484 | -65 | -60.004 | 0 | 0 | 0,20 | 11 | 4.0215 |  | 205349 | -0.00136 | 0.0053
2023-05-31 | 20:50 | MP | 5/31 nulling |  | sMATF LED pri | 20.9858 | 23.1351 | 20.6689 | 24.2679 | 24.437 | 23.414 | [5 empty] | 1,76 |  |  |  | 205034 |  |  | 814.7 | 592.9
2023-05-31 | 21:09 | MP | 5/31 nulling |  | sMATF mirror | 20.9915 | 23.1675 | 20.7163 | 24.2751 | 24.437 | 23.414 | [5 empty] | 0,20 |  |  |  | 210941 | -0.0147 | 0.0035
2023-05-31 | 20:02 | MP | 5/31 nulling |  | sMATF LED pri | 20.9732 | 23.1265 | 20.7163 | 24.2751 | 24.437 | 23.414 | [5 empty] | 1,76 |  |  |  | 200249 |  |  | 814.6 | 592.66
2023-05-31 | 21:18 | MP | 5/31 nulling |  | sMATF mirror | 21.0759 | 23.1874 | 20.7476 | 24.353 | 24.437 | 23.414 | [5 empty] | 0,20 |  |  |  | 211849 | -0.0049 | 0.0008
2023-05-31 | 21:16 | MP | 5/31 nulling |  | sMATF LED pri | 20.9915 | 23.1675 | 20.7476 | 24.353 | 24.437 | 23.414 | [5 empty] | 1,76 |  |  |  | 211651 |  |  | 814.6 | 592.96
2023-05-31 | 21:32 | MP | 5/31 nulling |  | sMATF mirror | 21.104 | 23.1926 | 20.7548 | 24.379 | 24.437 | 23.414 | [5 empty] | 0,20 |  |  |  | 213244 | -0.0042 | 0.0008
2023-05-31 | 21:31 | MP | 5/31 nulling |  | sMATF LED pri | 21.0759 | 23.1874 | 20.7548 | 24.379 | 24.437 | 23.414 | [5 empty] | 1,76 |  |  |  | 213101 |  |  | 814.07 | 592.9
2023-06-01 | 9:15 | MP | 6/1 initial (erroneous) nulling |  | sMATF mirror | 21.104 | 23.1926 | 20.7548 | 24.379 | 24.437 | 23.414 | [5 empty] | 0,20 | bad return |  |  | 091553 | -0.0063 | 0.006 |  |  | RM1&2 left Enabled overnight. Disabled at 09:00. Here's First measure
2023-06-01 | 9:00 | MP | 6/1 initial (erroneous) nulling |  | sMATF LED pri | 21.104 | 23.1926 | 20.7548 | 24.379 | 24.437 | 23.414 | -2147.484 | -65 | -60.004 | 0 | 0 | 1,76 |  |  |  | 090024 |  |  | 820.78 | 588.596 | First measure
2023-06-01 | 9:48 | MP | 6/1 initial (erroneous) nulling |  | sMATF LED pri | 21.1718 | 23.2484 | 20.8515 | 24.4391 | 24.437 | 23.414 | [5 empty] | 1,76 |  |  |  | 094858 |  |  | 814.05 | 593.88 | Pete's null pattern (adj RM1 LED, AutoCol meas, RM2 adj)
2023-06-01 | 9:52 | MP | 6/1 initial (erroneous) nulling |  | sMATF mirror | 21.1718 | 23.2484 | 20.8515 | 24.4391 | 24.437 | 23.414 | [5 empty] | 0,20 |  |  |  | 095247 | 0.0013 | -0.0031
2023-06-01 | 10:29 | MP | 6/1 initial (erroneous) nulling |  | sMATF LED pri | 21.54 | 23.2098 | 20.8016 | 24.4267 | 24.437 | 23.414 | [5 empty] | 1,76 |  |  |  | 102915 |  |  | 814.6 | 594.6
2023-06-01 | 10:31 | MP | 6/1 initial (erroneous) nulling |  | sMATF mirror | 21.54 | 23.2098 | 20.8016 | 24.4267 | 24.437 | 23.414 | [5 empty] | 0,20 |  |  |  | 103138 | 0.0069 | -0.0058
2023-06-01 | 10:42 | MP | 6/1 initial (erroneous) nulling |  | sMATF LED pri | 21.12 | 23.18 | 20.7497 | 24.3901 | 24.437 | 23.414 | [5 empty] | 1,76 |  |  |  | 104257 |  |  | 814.15 | 594.5
2023-06-01 | 10:45 | MP | 6/1 initial (erroneous) nulling |  | sMATF mirror | 21.12 | 23.18 | 20.7497 | 24.3901 | 24.437 | 23.414 | [5 empty] | 0,20 |  |  |  | 104529 | -0.0008 | 0.0008
2023-06-01 | 10:50 | MP | 6/1 initial (erroneous) nulling |  | sMATF LED pri | 21.12 | 23.18 | 20.7497 | 24.3901 | 24.437 | 23.414 | [5 empty] | 1,76 |  |  |  | 105020 |  |  | 814.51 | 593.95 | 10:51 RM1s disabled
2023-06-01 | 11:17 | MP | NULLING |  | sMATF LED pri | 21.1132 | 23.1748 | 20.7497 | 24.3901 | 24.437 | 23.414 | [5 empty] | 1,76 |  |  |  | 111733 |  |  | 807.68 | 601.9
2023-06-01 | 11:19 | MP | NULLING |  | sMATF mirror | 21.1132 | 23.1748 | 20.7497 | 24.3901 | 24.437 | 23.414 | -2147.484 | -65 | -60.004 | 0 | 0 | 0,20 |  |  |  | 111907 | 0.009 | -0.0076
2023-06-01 | 11:19 | MP | NULLING |  | sMATF LED pri | 21.0635 | 23.13 | 20.6817 | 24.3424 | 24.437 | 23.414 | [5 empty] | 1,76 |  |  |  | 111907 |  |  | 807.77 | 602.5
2023-06-01 | 11:13 | MP | NULLING |  | sMATF mirror | 21.0635 | 23.13 | 20.6817 | 24.3424 | 24.437 | 23.414 | [5 empty] | 0,20 |  |  |  | 111327 | 0.0037 | -0.002 |  |  | 11:41 RM1 disabled
2023-06-01 | 11:47 | MP | NULLING |  | sMATF LED pri | 21.0627 | 23.1288 | 20.6817 | 24.3424 | 24.437 | 23.414 | [5 empty] | 1,76 |  |  |  | 114731 |  |  | 807 | 602.4 | Immediate check after null showed drift, so removed RM1
2023-06-01 | 11:49 | MP | NULLING |  | sMATF mirror | 21.0627 | 23.1288 | 20.6817 | 24.3424 | 24.437 | 23.414 | [5 empty] | 0,20 |  |  |  | 114923 | 0.0041 | -0.0037
2023-06-01 | 11:50 | MP | NULLING |  | sMATF LED pri | 21.0627 | 23.1288 | 20.6817 | 24.3424 | 24.437 | 23.414 | [5 empty] | 1,76 |  |  |  | 115000 |  |  | 807.3 | 603.2 | 11:53 RM1s disabled
2023-06-01 | 11:51 | MP | SETTLING |  | sMATF LED pri | 21.0627 | 23.1288 | 20.6817 | 24.3424 | 24.437 | 23.414 | [5 empty] | 1,76 |  |  |  | 115100 |  |  | 809.8 | 598.3
2023-06-01 | 12:10 | MP | SETTLING |  | sMATF LED pri | 21.0627 | 23.1288 | 20.6817 | 24.3424 | 24.437 | 23.414 | [5 empty] | 1,76 |  |  |  | 121030 |  |  | 810.95 | 597.73
2023-06-01 | 12:12 | MP | SETTLING |  | sMATF mirror | 21.0627 | 23.1288 | 20.6817 | 24.3424 | 24.437 | 23.414 | [5 empty] | 0,20 |  |  |  | 121249 | -0.0015 | 0.0025
2023-06-01 | 12:19 | MP | SETTLING |  | sMATF LED pri | 21.0627 | 23.1288 | 20.6817 | 24.3424 | 24.437 | 23.414 | [5 empty] | 1,76 |  |  |  | 121926 |  |  | 812.6 | 598.16
2023-06-01 | 12:22 | MP | SETTLING |  | sMATF LED pri | 21.0627 | 23.1288 | 20.6817 | 24.3424 | 24.437 | 23.414 | [5 empty] | 1,76 |  |  |  | 122215 |  |  | 811.9 | 597.4
2023-06-01 | 12:25 | MP | SETTLING |  | sMATF LED pri | 21.0627 | 23.1288 | 20.6817 | 24.3424 | 24.437 | 23.414 | [5 empty] | 1,76 |  |  |  | 122544 |  |  | 811.99 | 597.39
2023-06-01 | 12:29 | MP | SETTLING |  | sMATF LED pri | 21.0627 | 23.1288 | 20.6817 | 24.3424 | 24.437 | 23.414 | [5 empty] | 1,76 |  |  |  | 122943 |  |  | 812.7 | 597.37
2023-06-01 | 12:35 | MP | SETTLING |  | sMATF LED pri | 21.0627 | 23.1288 | 20.6817 | 24.3424 | 24.437 | 23.414 | [5 empty] | 1,76 |  |  |  | 123508 |  |  | 812.75 | 597.3
2023-06-01 | 12:36 | MP | SETTLING |  | sMATF mirror | 21.0627 | 23.1288 | 20.6817 | 24.3424 | 24.437 | 23.414 | [5 empty] | 0,20 |  |  |  | 123644 | -0.0025 | 0.0032
2023-06-01 | 12:45 | MP | SETTLING |  | sMATF LED pri | 21.0627 | 23.1288 | 20.6817 | 24.3424 | 24.437 | 23.414 | [5 empty] | 1,76 |  |  |  | 124552 |  |  | 812.668 | 596.6
2023-06-01 | 12:50 | MP | SETTLING |  | sMATF LED pri | 21.0627 | 23.1288 | 20.6817 | 24.3424 | 24.437 | 23.414 | [5 empty] | 1,76 |  |  |  | 125000 |  |  | 812.82 | 596.65
2023-06-01 | 13:29 | MP | SETTLING |  | sMATF LED pri | 21.0627 | 23.1288 | 20.6817 | 24.3424 | 24.437 | 23.414 | [5 empty] | 1,76 |  |  |  | 132900 |  |  | 813.63 | 596.7
2023-06-01 | 13:31 | MP | SETTLING |  | sMATF LED pri | 21.0627 | 23.1288 | 20.6817 | 24.3424 | 24.437 | 23.414 | [5 empty] | 1,76 |  |  |  | 133132 |  |  | 812.93 | 596.7
2023-06-01 | 13:33 | MP | SETTLING |  | sMATF mirror | 21.0627 | 23.1288 | 20.6817 | 24.3424 | 24.425 | 23.422 | -2147.484 | -65 | -60.004 | 0 | 0 | 0,20 | 20 | 4.0215 |  | 133353 | -0.0022 | 0.0042
2023-06-01 | 14:48 | MP | 6/1 BA4 | Y | sMATF mirror | 21.0627 | 23.1288 | 20.6817 | 24.3424 | 24.425 | 23.422 | [5 empty] | 0,20 | 20 | 4.0215 |  | 144828 | -0.0018 | 0.0042 |  |  | Repeated sMATF after ASIST proc/git branch issues
2023-06-01 | 14:49 | MP | 6/1 BA4 | Y | sMATF LED pri | 21.0627 | 23.1288 | 20.6817 | 24.3424 | 24.425 | 23.422 | [5 empty] | 1,76 |  |  |  | 144957 |  |  | 813.64 | 597.54
2023-06-01 | 15:08 | MP | 6/1 BA4 | Y | PATB mirror | 21.0627 | 23.1288 | 20.6817 | 24.3424 | 24.425 | 23.422 | 101.513 | 125.725 | -0.004 | 0.014 | 0.0737 | 0,20 | 37 | 3.4722 |  | 150829 | -0.0011 | 0.0022 |  |  | no tuning done
2023-06-01 | 15:20 | MP | 6/1 BA4 | Y | PATB LED pri | 21.0627 | 23.1288 | 20.6817 | 24.3424 | 24.425 | 23.422 | 88.936 | 107.73 | 0 | 0.0143 | 0.0735 | 1130 |  | -6e-06 |  | 152008 | 0.0007 | -0.002 | 812.96 | 596.8 | no tuning done
2023-06-01 | 15:41 | MP | 6/1 BA4 |  | sMATF mirror | 21.0627 | 23.1288 | 20.6817 | 24.3424 | 24.425 | 23.422 | -2147.484 | -65 | -60.01 | -0.002 | 0.003 | 0,20 | 20 | 4.0215 |  | 154123 | -0.0015 | 0.0042
2023-06-01 | 15:44 | MP | 6/1 BA4 |  | sMATF LED pri | 21.0627 | 23.1288 | 20.6817 | 24.3424 | 24.425 | 23.422 | -2147.484 | -65 | -60.01 | -0.002 | 0.003 | 1,76 |  |  |  | 154402 |  |  | 813.66 | 597.52
2023-06-01 | 16:44 | MP | BA5 no-null |  | sMATF mirror | 21.0627 | 23.1288 | 20.6817 | 24.3424 | 24.465 | 23.472 | [5 empty] | 0,20 | 19 | 4.0221 |  | 164432 | 0.069 | 0.065 |  |  | 1022,808 initial >> measured at 897, 697
2023-06-01 | 16:54 | MP | BA5 no-null |  | sMATF LED pri | 21.0627 | 23.1288 | 20.6817 | 24.3424 | 24.465 | 23.472 | -2147.484 | -65 | -60.019 | -0.002 | 0.003 | 1,76 |  |  |  | 165425 |  |  | 812.927 | 594.985
2023-06-01 | 17:06 | MP | BA5 no-null |  | sMATF retro | 21.0627 | 23.1288 | 20.6817 | 24.3424 | 24.12 | 23.72 | [6 empty] | 31 | 4.0326 | [6 empty] | 17:29 RM3 disabled
2023-06-01 | 17:16 | MP | BA5 no-null |  | sMATF LED pri | 21.0627 | 23.1288 | 20.6817 | 24.3424 | 24.465 | 23.472 | [5 empty] | 1,76 |  |  |  | 171616 |  |  | 812.82 | 595.61 | pre tooling ball install
2023-06-01 | 17:31 | MP | BA5 no-null |  | sMATF LED pri | 21.0627 | 23.1288 | 20.6817 | 24.3424 | 24.465 | 23.472 | [5 empty] | 1,76 |  |  |  | 173105 |  |  | 812.62 | 594.7 | post-tooling ball install
2023-06-01 | 17:33 | MP | BA5 no-null |  | sMATF mirror | 21.0627 | 23.1288 | 20.6817 | 24.3424 | 24.465 | 23.472 | [5 empty] | 0,20 | 13 | 4.0221 |  | 173341 | 0.0683 | 0.0653
2023-06-01 | 17:55 | MP | BA5 no-null |  | PATB mirror | 21.0627 | 23.1288 | 20.6817 | 24.3424 | 24.425 | 23.422 | 101.513 | 125.725 | -0.004 | 0.014 | 0.0737 | 0,20 | 15 | 3.4722 |  | 175522 | -0.0015 | 0.0035
2023-06-01 | 18:08 | MP | BA5 no-null |  | PATB LED pri | 21.0627 | 23.1288 | 20.6817 | 24.3424 | 24.425 | 23.422 | 88.936 | 107.73 | -0.003 | 0.0143 | 0.0735 | 1132 |  |  |  | 180825 |  |  | 813.63 | 596.59
2023-06-01 | 18:27 | MP | BA5 no-null |  | sMATF mirror | 21.0627 | 23.1288 | 20.6817 | 24.3424 | 24.465 | 23.472 | [5 empty] | 0,20 | 8 | 4.0221 |  | 182702 | 0.0687 | 0.0643
2023-06-01 | 18:33 | MP | BA5 no-null |  | sMATF LED pri | 21.0627 | 23.1288 | 20.6817 | 24.3424 | 24.465 | 23.472 | [5 empty] | 1,76 |  |  |  | 183317 |  |  | 813.56 | 594.66
2023-06-01 | 19:21 | MP | NULLING |  | sMATF mirror | 20.312 | 23.7254 | 21.5236 | 23.6494 | 24.465 | 23.472 | [5 empty] | 0,20 |  |  |  | 192122 | -0.0129 | 0.0275
2023-06-01 |  | MP | NULLING |  | sMATF LED pri | 20.312 | 23.7254 | 21.5236 | 23.6494 | 24.465 | 23.472 | -2147.484 | -65 | -60.019 | -0.002 | 0.003 | 1,76 | [6 empty] | 807.85 | 601.81
2023-06-01 | 19:40 | MP | NULLING |  | sMATF mirror | 20.2972 | 23.7116 | 21.5236 | 23.6494 | 24.465 | 23.472 | [5 empty] | 0,20 |  |  |  | 194049 | 0.0173 | 0.008
2023-06-01 | 19:45 | MP | NULLING |  | sMATF LED pri | 20.2972 | 23.7116 | 21.5236 | 23.6494 | 24.465 | 23.472 | [5 empty] | 1,76 |  |  |  | 194548 |  |  | 807.67 | 601.76
2023-06-01 | 19:48 | MP | NULLING |  | sMATF mirror | 20.302 | 23.7142 | 21.5236 | 23.6494 | 24.465 | 23.472 | [5 empty] | 0,20 |  |  |  | 194823 | 0.0103 | 0.0049
2023-06-01 | 19:52 | MP | NULLING |  | sMATF LED pri | 20.184 | 23.7764 | 21.6113 | 23.5402 | 24.465 | 23.472 | [5 empty] | 1,76 |  |  |  | 195200 |  |  | 807.793 | 601.774
2023-06-01 | 19:58 | MP | NULLING |  | sMATF mirror | 20.184 | 23.7764 | 21.6113 | 23.5402 | 24.465 | 23.472 | [5 empty] | 0,20 |  |  |  | 195807 | -0.001 | -0.0013
2023-06-01 | 20:00 | MP | NULLING |  | sMATF LED pri | 20.184 | 23.7764 | 21.6113 | 23.5402 | 24.465 | 23.472 | [5 empty] | 1,76 |  |  |  | 200030 |  |  | 807.6 | 601.64 | 8:03:00 PM RMs1&2 disabled
2023-06-01 |  | MP | SETTLING |  |  | 20.184 | 23.7764 | 21.6113 | 23.5402 | 24.465 | 23.472 | [14 empty] | live camera monitoring LED settling until 8:21 PM
2023-06-01 | 20:23 | MP | SETTLING |  | sMATF LED pri | 20.184 | 23.7764 | 21.6113 | 23.5402 | 24.465 | 23.472 | [5 empty] | 1,76 |  |  |  | 202314 |  |  | 809.685 | 601.59 | 8:21 PM turned off camera
2023-06-01 | 20:26 | MP | SETTLING |  | sMATF mirror | 20.184 | 23.7764 | 21.6113 | 23.5402 | 24.465 | 23.472 | [5 empty] | 0,20 |  |  |  | 202618 | -0.0008 | 0.0008
2023-06-01 | 20:39 | MP | SETTLING |  | sMATF mirror | 20.184 | 23.7764 | 21.6113 | 23.5402 | 24.465 | 23.472 | [5 empty] | 0,20 |  |  |  | 203930 | -0.0008 | 0.0015
2023-06-01 | 20:41 | MP | SETTLING |  | sMATF LED pri | 20.184 | 23.7764 | 21.6113 | 23.5402 | 24.465 | 23.472 | [5 empty] | 1,76 |  |  |  | 204133 |  |  | 810.57 | 600.64 | Null the PATB mirror/LED to the post-settling sMATF values.
2023-06-01 | 20:49 | MP | SETTLING |  | sMATF LED pri | 20.184 | 23.7764 | 21.6113 | 23.5402 | 24.465 | 23.472 | [5 empty] | 1,76 |  |  |  | 204951 |  |  | 810.6 | 600.65 | Text Evan when you start the move back to sMATF
2023-06-01 | 20:52 | MP | SETTLING |  | sMATF mirror | 20.184 | 23.7764 | 21.6113 | 23.5402 | 24.465 | 23.472 | [5 empty] | 0,20 |  |  |  | 205222 | -0.0008 | 0.0021
2023-06-01 | 20:59 | MP | SETTLING |  | sMATF mirror | 20.184 | 23.7764 | 21.6113 | 23.5402 | 24.465 | 23.472 | [5 empty] | 0,20 |  |  |  | 205907 | -0.0004 | 0.0022
2023-06-01 | 21:01 | MP | SETTLING |  | sMATF LED pri | 20.184 | 23.7764 | 21.6113 | 23.5402 | 24.465 | 23.472 | [5 empty] | 1,76 |  |  |  | 210157 |  |  | 810.65 | 600.623
2023-06-01 | 21:09 | MP | SETTLING |  | sMATF LED pri | 20.184 | 23.7764 | 21.6113 | 23.5402 | 24.465 | 23.472 | [5 empty] | 1,76 |  |  |  | 210912 |  |  | 810.65 | 600.64
2023-06-01 | 21:11 | MP | SETTLING |  | sMATF mirror | 20.184 | 23.7764 | 21.6113 | 23.5402 | 24.465 | 23.472 | [5 empty] | 0,20 |  |  |  | 211140 | -0.0004 | 0.0022
2023-06-01 | 21:25 | MP | SETTLING |  | sMATF mirror | 20.184 | 23.7764 | 21.6113 | 23.5402 | 24.465 | 23.472 | [5 empty] | 0,20 |  |  |  | 212535 | -0.0004 | 0.0022
2023-06-01 | 21:26 | MP | SETTLING |  | sMATF LED pri | 20.184 | 23.7764 | 21.6113 | 23.5402 | 24.465 | 23.472 | [5 empty] | 1,76 |  |  |  | 212635 |  |  | 810.675 | 600.71
2023-06-01 | 21:37 | MP | SETTLING |  | sMATF LED pri | 20.184 | 23.7764 | 21.6113 | 23.5402 | 24.465 | 23.472 | [5 empty] | 1,76 |  |  |  | 213758 |  |  | 810.66 | 600.78
2023-06-01 | 21:40 | MP | SETTLING |  | sMATF mirror | 20.184 | 23.7764 | 21.6113 | 23.5402 | 24.465 | 23.472 | [5 empty] | 0,20 |  |  |  | 214012 | -0.0001 | 0.0022
2023-06-01 | 22:14 | MP | BA5 with null | Y | sMATF mirror | 20.184 | 23.7764 | 21.6113 | 23.5402 | 24.44 | 23.42 | -2147.484 | -65 | -60.019 | -0.002 | 0.003 | 0,20 | 20 | 4.0219 |  | 2214950 | -0.0004 | 0.0022
2023-06-01 | 22:00 | MP | BA5 with null | Y | sMATF LED pri | 20.184 | 23.7764 | 21.6113 | 23.5402 | 24.44 | 23.42 | -2147.484 | -65 | -60.019 | -0.002 | 0.003 | 1,76 |  | -5.4e-05 |  | 220010 |  |  | 810.66 | 601.65
2023-06-01 | 22:19 | MP | BA5 with null | Y | PATB mirror | 20.184 | 23.7764 | 21.6113 | 23.5402 | 24.44 | 23.42 | 101.513 | 125.728 | -0.004 | 0.1076 | 0.0876 | 0,20 | 24 | 3.4722 |  | 221911 | -0.0008 | 0.0032 |  |  | 22:34 RM3 disabled
2023-06-01 | 22:54 | MP | BA5 with null | Y | PATB LED pri | 20.184 | 23.7764 | 21.6113 | 23.5402 | 24.44 | 23.42 | 88.893 | 108.635 | -0.004 | 0.1078 | 0.0877 | 1132 |  | -3.9e-05 |  | 225454 |  |  | 810.69 | 601.72
2023-06-01 | 23:15 | MP | BA5 with null |  | sMATF mirror | 20.184 | 23.7764 | 21.6113 | 23.5402 | 24.44 | 23.42 | -2147.484 | -65 | -60.019 | -0.002 | 0.003 | 0,20 | 21 | 4.0219 |  | 231549 | 0.0003 | 0.0022 |  | 0.07
2023-06-01 | 23:18 | MP | BA5 with null |  | sMATF LED pri | 20.184 | 23.7764 | 21.6113 | 23.5402 | 24.44 | 23.42 | -2147.484 | -65 | -60.019 | -0.002 | 0.003 | 1,76 |  | 5e-06 |  | 231849 |  |  | 810.68 | 600.8
2023-06-03 | 9:44 | MP | NULLING |  | sMATF mirror | 20.184 | 23.7764 | 21.6113 | 23.5402 | 24.42 | 23.41 | -2147.484 | -65 | -60.019 | -0.002 | 0.003 | 0,20 | 16 | 4.0219 |  | 094435 | -0.0028 | -0.0381 |  |  | Homed RMs. GO to previous positions. RM1a jogged back to same number dif position. RM1b jogged. RM2s not moved.
2023-06-03 | 9:53 | MP | NULLED |  | sMATF LED pri | 20.184 | 23.798 | 21.6113 | 23.5402 | 24.42 | 23.41 | [5 empty] | 1,76 |  |  |  | 095307 |  |  | 807.596 | 601.739 | (The RM2 mirrors are not positioned like I normally do)
2023-06-03 | 10:06 | MP | NULLED |  | sMATF mirror | 20.184 | 23.798 | 21.6113 | 23.5402 | 24.42 | 23.41 | [5 empty] | 0,20 |  |  |  | 100658 | -0.0003 | -0.002 |  |  | 10:08: All RMs off, starting settle
2023-06-03 | 10:10 | MP | SETTLING |  | sMATF LED pri | 20.184 | 23.798 | 21.6113 | 23.5402 | 24.42 | 23.41 | [5 empty] | 1,76 |  |  |  | 101008 |  |  | 807.778 | 600.667
2023-06-03 | 10:28 | MP | SETTLING |  | sMATF mirror | 20.184 | 23.798 | 21.6113 | 23.5402 | 24.42 | 23.41 | [5 empty] | 0,20 |  |  |  | 102851 | -0.0025 | 0.0001 | 0.182 | -1.072
2023-06-03 | 10:26 | MP | SETTLING |  | sMATF LED pri | 20.184 | 23.798 | 21.6113 | 23.5402 | 24.42 | 23.41 | [5 empty] | 1,76 |  |  |  | 102645 |  |  | 808.74 | 599.64
2023-06-03 | 10:42 | MP | SETTLING |  | sMATF mirror | 20.184 | 23.798 | 21.6113 | 23.5402 | 24.42 | 23.41 | [5 empty] | 0,20 |  |  |  | 104236 | -0.0028 | 0.0004 | 0 | 0
2023-06-03 | 10:43 | MP | SETTLING |  | sMATF LED pri | 20.184 | 23.798 | 21.6113 | 23.5402 | 24.42 | 23.41 | [5 empty] | 1,76 |  |  |  | 104348 |  |  | 808.83 | 598.8
2023-06-03 | 10:53 | MP | SETTLING |  | sMATF mirror | 20.184 | 23.798 | 21.6113 | 23.5402 | 24.42 | 23.41 | [5 empty] | 0,20 |  |  |  | 105333 | -0.0028 | 0.0004 | 0.09 | -0.84
2023-06-03 | 10:52 | MP | SETTLING |  | sMATF LED pri | 20.184 | 23.798 | 21.6113 | 23.5402 | 24.42 | 23.41 | [5 empty] | 1,76 |  |  |  | 105215 |  |  | 808.793 | 599.56
2023-06-03 | 10:53 | MP | BA5 repeatability | Y | sMATF mirror | 20.184 | 23.798 | 21.6113 | 23.5402 | 24.435 | 23.41 | -2147.484 | -65 | -60.019 | -0.002 | 0.003 | 0,20 | 15 | 4.0219 |  | 105333 | -0.0032 | 0.0004 | 0.053 | -0.08
2023-06-03 | 11:24 | MP | BA5 repeatability | Y | sMATF LED pri | 20.184 | 23.798 | 21.6113 | 23.5402 | 24.435 | 23.41 | -2147.484 | -65 | -60.019 | -0.002 | 0.003 | 1,76 |  |  |  | 112448 |  |  | 808.8 | 599.6
2023-06-03 | 11:24 | MP | BA5 repeatability | Y | PATB mirror | 20.184 | 23.798 | 21.6113 | 23.5402 | 24.435 | 23.41 | 101.513 | 125.728 | 0.001 | 0.1076 | 0.0877 | 0,20 | 31 | 3.4721 |  | 112448 | -0.0039 | 0.0008 | 0.06 | -0.04
2023-06-03 | 12:14 | MP | BA5 repeatability | Y | PATB LED pri | 20.184 | 23.798 | 21.6113 | 23.5402 | 24.435 | 23.41 | 88.893 | 108.635 | -0.004 | 0.1078 | 0.0877 | 1132 |  |  |  | 121448 | -0.0007 | 0.0004 | 808.8 | 599.9
2023-06-03 | 12:51 | MP | BA5 repeatability | Y | sMATF mirror | 20.184 | 23.798 | 21.6113 | 23.5402 | 24.435 | 23.41 | -2147.484 | -65 | -60.019 | -0.002 | 0.003 | 0,20 | 17 | 4.0219 |  | 125142 | -0.0028 | 0.0004 | 0 | 0.3
2023-06-03 | 12:47 | MP | BA5 repeatability | Y | sMATF LED pri | 20.184 | 23.798 | 21.6113 | 23.5402 | 24.435 | 23.41 | -2147.484 | -65 | -60.019 | -0.002 | 0.003 | 1,76 |  |  |  | 124731 |  |  | 808.8 | 598.9
2023-06-03 | 15:03 | MP | NULLED |  | sMATF mirror | 20.1892 | 23.7984 | 21.6113 | 23.5402 | 24.425 | 23.415 | -2147.484 | -65 | -60.019 | -0.002 | 0.003 | 0,20 | 15 | 4.0219 |  | 150304 | 0.002 | -0.0024 | 0 | -0.7 | lost control of RM1a 13:41, Restart FARM
2023-06-03 | 15:01 | MP | NULLED |  | sMATF LED pri | 20.1892 | 23.7984 | 21.6113 | 23.5402 | 24.425 | 23.415 | [5 empty] | 1,76 |  |  |  | 150114 |  |  | 807.57 | 601.66 | 1510: All RMS off, starting settle
2023-06-03 | 15:29 | MP | SETTLING |  | sMATF mirror | 20.1892 | 23.7984 | 21.6113 | 23.5402 | 24.425 | 23.415 | [5 empty] | 0,20 |  |  |  | 152902 | 0.0003 | 0.0001
2023-06-03 | 15:30 | MP | SETTLING |  | sMATF LED pri | 20.1892 | 23.7984 | 21.6113 | 23.5402 | 24.425 | 23.415 | [5 empty] | 1,76 |  |  |  | 153022 |  |  | 807.67 | 601.59
2023-06-03 | 15:41 | MP | SETTLING |  | sMATF mirror | 20.1892 | 23.7984 | 21.6113 | 23.5402 | 24.425 | 23.415 | [5 empty] | 0,20 |  |  |  | 154107 | 0.0003 | 0.0001
2023-06-03 | 15:39 | MP | SETTLING |  | sMATF LED pri | 20.1892 | 23.7984 | 21.6113 | 23.5402 | 24.425 | 23.415 | [5 empty] | 1,76 |  |  |  | 153958 |  |  | 808.58 | 601.62
2023-06-03 | 15:47 | MP | SETTLING |  | sMATF mirror | 20.1892 | 23.7984 | 21.6113 | 23.5402 | 24.425 | 23.415 | [5 empty] | 0,20 |  |  |  | 154758 | 0.0003 | 0.0001
2023-06-03 | 15:48 | MP | SETTLING |  | sMATF LED pri | 20.1892 | 23.7984 | 21.6113 | 23.5402 | 24.425 | 23.415 | [5 empty] | 1,76 |  |  |  | 154846 |  |  | 808.54 | 601.6
2023-06-03 | 16:02 | MP | BA5 repeatability | Y | sMATF mirror | 20.1892 | 23.7984 | 21.6113 | 23.5402 | 24.425 | 23.415 | -2147.484 | -65 | -60.019 | -0.002 | 0.003 | 0,20 | 19 | 4.0219 |  | 160204 | 0.0003 | 0.0001
2023-06-03 | 16:03 | MP | BA5 repeatability | Y | sMATF LED pri | 20.1892 | 23.7984 | 21.6113 | 23.5402 | 24.425 | 23.415 | -2147.484 | -65 | -60.019 | -0.002 | 0.003 | 1,76 |  |  |  | 160312 |  |  | 808.56 | 602.455
2023-06-03 | 16:26 | MP | BA5 repeatability | Y | PATB mirror | 20.1892 | 23.7984 | 21.6113 | 23.5402 | 24.425 | 23.415 | 101.513 | 125.728 | 0.001 | 0.1072 | 0.0852 | 0,20 |  | oops |  | 162637 | 0.0003 | 0.0006
2023-06-03 | 16:32 | MP | BA5 repeatability | Y | PATB LED pri | 20.1892 | 23.7984 | 21.6113 | 23.5402 | 24.425 | 23.415 | 88.893 | 108.635 | -0.004 | 0.1072 | 0.0852 | 1132 |  |  |  | 163201 |  |  | 808.64 | 602.5
2023-06-03 | 16:48 | MP | BA5 repeatability | Y | sMATF mirror | 20.1892 | 23.7984 | 21.6113 | 23.5402 | 24.425 | 23.415 | -2147.484 | -65 | -60.019 | -0.002 | 0.003 | 0,20 | 19 | 4.0219 |  | 164819 | 0.0003 | 0.0001 | 0.08 | 0.045
2023-06-03 | 16:51 | MP | BA5 repeatability | Y | sMATF LED pri | 20.1892 | 23.7984 | 21.6113 | 23.5402 | 24.425 | 23.415 | -2147.484 | -65 | -60.019 | -0.002 | 0.003 | 1,76 |  |  |  | 165121 |  |  | 808.721 | 601.71
2023-06-03 | 18:24 | MP | NULLED |  | sMATF mirror | 20.1683 | 23.7452 | 21.5543 | 23.5222 | 24.425 | 23.415 | [5 empty] | 0,20 |  |  |  | 182458 | 0.0024 | -0.0027 |  |  | Rehomed mirrors again (twice this time cuz didn't seem to come back to same position).
2023-06-03 | 18:23 | MP | NULLED |  | sMATF LED pri | 20.2011 | 23.7877 | 21.6151 | 23.5519 | 24.425 | 23.415 | -2147.484 | -65 | -60.019 | -0.002 | 0.003 | 1,76 |  |  |  | 182301 |  |  | 807.66 | 601.76 | This previous two nulls were very easy compared to this one
2023-06-03 | 18:00 | MP | SETTLING |  | sMATF LED pri | 20.2011 | 23.7877 | 21.6151 | 23.5519 | 24.425 | 23.415 | [5 empty] | 1,76 |  |  |  | 180010 |  |  | 806.5 | 602.7 | RMs are off
2023-06-03 | 18:23 | MP | SETTLING |  | sMATF mirror | 20.2011 | 23.7877 | 21.6151 | 23.5519 | 24.425 | 23.415 | [5 empty] | 0,20 |  |  |  | 182301 | -0.0028 | 0.0032
2023-06-03 | 18:43 | MP | SETTLING |  | sMATF LED pri | 20.2011 | 23.7877 | 21.6151 | 23.5519 | 24.425 | 23.415 | [5 empty] | 1,76 |  |  |  | 184314 |  |  | 809.83 | 599.61
2023-06-03 | 18:41 | MP | SETTLING |  | sMATF mirror | 20.2011 | 23.7877 | 21.6151 | 23.5519 | 24.425 | 23.415 | [5 empty] | 0,20 |  |  |  | 184150 | -0.0011 | 0.0011 |  |  | SMC functional is running in parallel
2023-06-03 | 18:52 | MP | SETTLING |  | sMATF LED pri | 20.2011 | 23.7877 | 21.6151 | 23.5519 | 24.425 | 23.415 | [5 empty] | 1,76 |  |  |  | 185243 |  |  | 810.6 | 598.83
2023-06-03 | 18:54 | MP | SETTLING |  | sMATF mirror | 20.2011 | 23.7877 | 21.6151 | 23.5519 | 24.425 | 23.415 | [5 empty] | 0,20 |  |  |  | 185419 | -0.0015 | 0.0015 | 0 | 0
2023-06-03 | 19:04 | MP | SETTLING |  | sMATF LED pri | 20.2011 | 23.7877 | 21.6151 | 23.5519 | 24.425 | 23.415 | [5 empty] | 1,76 |  |  |  | 190413 |  |  | 810.65 | 598.78
2023-06-03 | 19:04 | MP | SETTLING |  | sMATF mirror | 20.2011 | 23.7877 | 21.6151 | 23.5519 | 24.425 | 23.415 | [5 empty] | 0,20 |  |  |  | 190413 | -0.0015 | 0.0015
2023-06-03 | 19:13 | MP | SETTLING |  | sMATF LED pri | 20.2011 | 23.7877 | 21.6151 | 23.5519 | 24.425 | 23.415 | [5 empty] | 1,76 |  |  |  | 191342 |  |  | 810.61 | 599.5 | We've reached stability here
2023-06-03 | 19:15 | MP | SETTLING |  | sMATF mirror | 20.2011 | 23.7877 | 21.6151 | 23.5519 | 24.425 | 23.415 | [5 empty] | 0,20 |  |  |  | 191517 | -0.0015 | 0.0022
2023-06-03 | 19:31 | MP | BA5 repeatability | Y | sMATF mirror | 20.2011 | 23.7877 | 21.6151 | 23.5519 | 24.44 | 23.416 | -2147.484 | -65 | -60.019 | -0.002 | 0.003 | 0,20 | 19 | 4.0219 |  | 193110 | -0.0018 | 0.0022
2023-06-03 | 19:33 | MP | BA5 repeatability | Y | sMATF LED pri | 20.2011 | 23.7877 | 21.6151 | 23.5519 | 24.44 | 23.416 | -2147.484 | -65 | -60.019 | -0.002 | 0.003 | 1,76 |  |  |  | 193300 |  |  | 810.66 | 598.79
2023-06-03 | 20:06 | MP | BA5 repeatability | Y | PATB mirror | 20.2011 | 23.7877 | 21.6151 | 23.5519 | 24.44 | 23.416 | 101.513 | 125.728 | 0.001 | 0.1072 | 0.0852 | 0,22 | 50 | 3.4722 |  | 200641 | -0.0015 | 0.0014 | 0.06 | -0.04 | changed RM3 to 24.426 and got 3.472159
2023-06-03 | 20:24 | MP | BA5 repeatability | Y | PATB LED pri | 20.2011 | 23.7877 | 21.6151 | 23.5519 | 24.44 | 23.416 | 88.893 | 108.635 | -0.004 | 0.1072 | 0.0852 | 1133 |  |  |  | 202405 |  |  | 809.8 | 599.68
2023-06-03 | 20:40 | MP | BA5 repeatability | Y | sMATF mirror | 20.2011 | 23.7877 | 21.6151 | 23.5519 | 24.44 | 23.416 | -2147.484 | -65 | -60.019 | -0.002 | 0.003 | 0,20 | 35 | 4.0219 |  | 204042 | -0.0015 | 0.0022 | -0.8 | 0.85
2023-06-03 | 20:45 | MP | BA5 repeatability | Y | sMATF LED pri | 20.2011 | 23.7877 | 21.6151 | 23.5519 | 24.44 | 23.416 | -2147.484 | -65 | -60.019 | -0.002 | 0.003 | 1,76 |  |  |  | 204520 |  |  | 810.64 | 598.799
2023-06-05 | 14:03 | MP | NULLED |  | sMATF mirror | 20.3263 | 23.6265 | 21.5619 | 23.6801 | 24.44 | 23.416 | [5 empty] | 0,20 | 35 | 4.0272 |  | 140354 | -0.0011 | 0.0011 |  |  | 14:02 RMs OFF
2023-06-05 | 14:01 | MP | NULLED |  | sMATF LED pri | 20.3323 | 23.6257 | 21.5334 | 23.6674 | 24.44 | 23.416 | -2147.484 | -65 | -60.019 | -0.002 | 0.003 | 1,76 |  |  |  | 140147 |  |  | 807.67 | 601.56
2023-06-05 | 14:11 | MP | SETTLING |  | sMATF mirror | 20.3323 | 23.6257 | 21.5334 | 23.6674 | 24.44 | 23.416 | [5 empty] | 0,20 |  |  |  | 141127 | -0.0025 | 0.0028
2023-06-05 | 14:12 | MP | SETTLING |  | sMATF LED pri | 20.3323 | 23.6257 | 21.5334 | 23.6674 | 24.44 | 23.416 | [5 empty] | 1,76 |  |  |  | 141226 |  |  | 809.775 | 598.696
2023-06-05 | 14:24 | MP | SETTLING |  | sMATF mirror | 20.3323 | 23.6257 | 21.5334 | 23.6674 | 24.44 | 23.416 | [5 empty] | 0,20 |  |  |  | 142432 | -0.0028 | 0.0032 | 0 | 0
2023-06-05 | 14:22 | MP | SETTLING |  | sMATF LED pri | 20.3323 | 23.6257 | 21.5334 | 23.6674 | 24.44 | 23.416 | [5 empty] | 1,76 |  |  |  | 142241 |  |  | 809.9 | 598.6
2023-06-05 | 14:32 | MP | SETTLING |  | sMATF mirror | 20.3323 | 23.6257 | 21.5334 | 23.6674 | 24.44 | 23.416 | [5 empty] | 0,20 |  |  |  | 143245 | -0.0028 | 0.0035 | 0.125 | -0.096
2023-06-05 | 14:33 | MP | SETTLING |  | sMATF LED pri | 20.3323 | 23.6257 | 21.5334 | 23.6674 | 24.44 | 23.416 | [5 empty] | 1,76 |  |  |  | 143337 |  |  | 810.68 | 597.796
2023-06-05 | 14:56 | MP | SETTLING |  | sMATF mirror | 20.3323 | 23.6257 | 21.5334 | 23.6674 | 24.44 | 23.416 | [5 empty] | 0,20 |  |  |  | 145626 | -0.0028 | 0.0035 | 0.905 | -0.9
2023-06-05 | 14:43 | MP | SETTLING |  | sMATF LED pri | 20.3323 | 23.6257 | 21.5334 | 23.6674 | 24.44 | 23.416 | [5 empty] | 1,76 |  |  |  | 144332 |  |  | 810.7 | 598.6
2023-06-05 | 15:01 | MP | SETTLING |  | sMATF mirror | 20.3323 | 23.6257 | 21.5334 | 23.6674 | 24.44 | 23.416 | [5 empty] | 0,20 |  |  |  | 150106 | -0.0028 | 0.0035 | 0.925 | -0.096
2023-06-05 | 15:12 | MP | SETTLING |  | sMATF LED pri | 20.3323 | 23.6257 | 21.5334 | 23.6674 | 24.44 | 23.416 | [5 empty] | 1,76 |  |  |  | 151200 |  |  | 810.66 | 598.607
2023-06-05 | 15:13 | MP | SETTLING |  | sMATF mirror | 20.3323 | 23.6257 | 21.5334 | 23.6674 | 24.44 | 23.416 | [5 empty] | 0,20 |  |  |  | 151302 | -0.0028 | 0.0035 | 0.885 | -0.089
2023-06-05 | 15:22 | MP | BA6 | Y | sMATF mirror | 20.3323 | 23.6257 | 21.5334 | 23.6674 | 24.44 | 23.416 | -2147.484 | -65 | -60.019 | -0.002 | 0.003 | 0,20 | 19 | 4.0217 |  | 152231 | -0.0028 | 0.0035
2023-06-05 | 15:20 | MP | BA6 | Y | sMATF LED pri | 20.3323 | 23.6257 | 21.5334 | 23.6674 | 24.44 | 23.416 | -2147.484 | -65 | -60.019 | -0.002 | 0.003 | 1,76 |  |  |  | 152038 |  |  | 810.685 | 598.598
2023-06-05 | 15:31 | MP | BA6 | Y | sMATF retro | 20.3323 | 23.6257 | 21.5334 | 23.6674 | 24.153 | 23.75 | -2147.484 | -65 | -60.019 | -0.002 | 0.003 | 0,20 | 15 | 4.0323
2023-06-05 | 15:33 | MP | BA6 | Y | CHIPS retro | 20.3323 | 23.6257 | 21.5334 | 23.6674 | 24.63 | 23.18 | -2147.484 | -65 | -60.019 | -0.002 | 0.003 | 0,20 | 7 | 4.7633 | [6 empty] | 15:33 RM3 off
2023-06-05 | 15:56 | MP | BA6 | Y | PATB mirror | 20.3323 | 23.6257 | 21.5334 | 23.6674 | 24.44 | 23.416 | 101.513 | 125.728 | 0.001 | 0.1058 | 0.0867 | 0,20 | 25 | 3.4722 |  | 155645 | -0.0025 | 0.0028
2023-06-05 | 16:14 | MP | BA6 | Y | PATB LED pri | 20.3323 | 23.6257 | 21.5334 | 23.6674 | 24.44 | 23.416 | 88.679 | 108.149 | -0.002 | 0.1058 | 0.0868 | 1131 |  |  |  | 161420 |  |  | 810.78 | 597.963
2023-06-05 | 16:21 | MP | BA6 | Y | PATB retro pri | 20.3323 | 23.6257 | 21.5334 | 23.6674 | 24.44 | 23.416 | 105.901 | 99.041 | -0.001 | 0.1061 | 0.0868 | 0,20 | 31 | 3.4586
2023-06-05 | 16:43 | MP | BA6 | Y | PATA mirror | 20.3323 | 23.6257 | 21.5334 | 23.6674 | 24.44 | 23.416 | -76.119 | 125.7 | -0.002 | 0.00816 | 0.0782 | 0,20 | 14 | 3.4722 |  | 164308 | -0.0035 | 0.0039
2023-06-05 | 16:53 | MP | BA6 | Y | PATA LED pri | 20.3323 | 23.6257 | 21.5334 | 23.6674 | 24.44 | 23.416 | -60.424 | 108.061 | -0.003 | 0.00816 | 0.0782 | 1128 |  |  |  | 165308 |  |  | 810.586 | 598.856
2023-06-05 | 16:58 | MP | BA6 | Y | PATA retro pri | 20.3323 | 23.6257 | 21.5334 | 23.6674 | 24.44 | 23.416 | -72.08 | 98.978 | 0 | 0.00816 | 0.0782 | 0,20 | 8 | 3.4586
2023-06-05 | 17:12 | MP | BA6 |  | sMATF mirror | 20.3323 | 23.6257 | 21.5334 | 23.6674 | 24.44 | 23.416 | -2147.484 | -65 | -60.019 | -0.002 | 0.003 | 0,20 | 11 | 4.0217 |  | 171205 | -0.0028 | 0.0039
2023-06-05 | 17:14 | MP | BA6 |  | sMATF LED pri | 20.3323 | 23.6257 | 21.5334 | 23.6674 | 24.44 | 23.416 | -2147.484 | -65 | -60.019 | -0.002 | 0.003 | 1,76 |  |  |  | 171448 | 0 | 0.0004 | 810.754 | 598.636
2023-06-05 | 17:18 | MP | BA6 |  | sMATF retro | 20.3323 | 23.6257 | 21.5334 | 23.6674 | 24.153 | 23.75 | -2147.484 | -65 | -60.019 | -0.002 | 0.003 | 0,20 | 11 | 4.0322 |  |  |  |  | 0.069 | 0.038
2023-06-05 | 17:20 | MP | BA6 |  | CHIPS retro | 20.3323 | 23.6257 | 21.5334 | 23.6674 | 24.63 | 23.18 | -2147.484 | -65 | -60.019 | -0.002 | 0.003 | 0,20 | 7 | 4.7633
2023-06-07 | 21:36 | CD | 3832 Event PR-009 continuation #2 |  |  | 20.3323 | 23.6257 | 20.93 | 23.37 | 24.153 | 23.75 | 525 | -400 | 0 | 0 | 0 | 1,76 |  |  |  | 212831 | 0.0152 | -0.1406 | 1251.6 | 175.5 | AC LEDs on sMATF primary retro and flat mirror pinhole in view (moved RM2 to put AC LEDs on retro)
2023-06-05 | 21:41 | CD | 3832 Event PR-009 continuation #2 |  |  | 20.3323 | 23.6257 | 20.93 | 23.37 | 24.153 | 23.75 | 525 | -400 | 0 | 0 | 0 | 0,20 |  |  |  | 214116 |  |  |  |  | tried to go to infinity focus but code bits are too blurry
2023-06-05 | 21:45 | CD | 3832 Event PR-009 continuation #2 |  |  | 20.3323 | 23.6257 | 20.68 | 23.37 | 24.153 | 23.75 | 525 | -400 | 0 | 0 | 0 | 0,20 |  |  |  | 214532 |  |  |  |  | maybe the code bits are better centered in the camera? is that even desirable? Still can't decode, though
2023-06-13 | 10:10 | MW | Power up ADM, Farm on with all RM disabled
2023-06-13 | 16:31 | MW | Take dark image and enable dark subtract | [17 empty] | 62927
2023-06-13 | 15:58 | MW | Home and reposition all RMs to row 308 |  |  | 20.3323 | 23.6257 | 21.5334 | 23.6674 | 24.44 | 23.416
2023-06-13 | 17:10 | MW | Null | y | sMATF LED pri | 20.8735 | 23.811 | 21.8164 | 24.149 | 24.44 | 23.42 | [5 empty] | 1,76 |  |  |  | 70704 |  |  | 807.362 | 602.373
2023-06-13 | 17:12 | MW | Nulled | y | sMATF mirror | 20.8735 | 23.811 | 21.8164 | 24.149 | 24.44 | 23.42 | [5 empty] | 0,16 |  |  |  | 71036 | -0.0035 | 0.0032
2023-06-14 | 10:11 | MW | ADM power left on overnight, Farm on with all RMs disabled. |  |  | 20.8735 | 23.811 | 21.8164 | 24.149
2023-06-14 | 11:39 | MW | BA7 |  | sMATF mirror | 20.8735 | 23.811 | 21.8164 | 24.149 | 24.434 | 23.418 | -2147 | -65 | 0 | 0 | 0 | 0,16 | 8 | 4.0239 |  | 13719 | -0.0063 | 0.0091 |  |  | Leica beam centroid is ~796/581 at 0,16 focus, standard deviation ~4um
2023-06-14 | 12:44 | MW | BA7 |  | sMATF LED pri | 20.8735 | 23.811 | 21.8164 | 24.149 | 24.434 | 23.418 | -2147 | -65 | 0 | 0 | 0 | 1,76 |  |  |  | 24229 |  |  | 814.584 | 597.53
2023-06-14 | 10:12 | MW | BA7 | Y | PATB mirror | 20.8735 | 23.811 | 21.8164 | 24.149 | 24.439 | 23.413 | 101.516 | 125.735 | 0.003 | 0.1085 | 0.0885 | 0,16 | 37 | 3.4744 |  | 34125 | -0.0056 | 0.0098 |  |  | Std dev~6um
2023-06-14 | 18:22 | MW | BA7 | Y | PATB LED pri | 20.8735 | 23.811 | 21.8164 | 24.149 | 24.439 | 23.413 | 88.753 | 108.118 | 0.003 | 0.1086 | 0.0885 | 1135 |  |  |  | 42227 |  |  | 814.804 | 597.623
2023-06-14 | 14:56 | MW | BA7 |  | sMATF mirror | 20.8735 | 23.811 | 21.8164 | 24.149 | 24.439 | 23.41 | -2147 | -65 | 0 | 0 | 0 | 0,16 | 8 | 4.024 |  | 44758 | -0.007 | 0.0077 |  |  | Std dev ~7um
2023-06-14 | 14:58 | MW | BA7 |  | sMATF LED pri | 20.8735 | 23.811 | 21.8164 | 24.149 | 24.439 | 23.413 | -2147 | -65 | 0 | 0 | 0 | 1,76 |  |  |  | 14513 |  |  | 813.75 | 598.131
2023-06-14 | - | EB | BA7 | Y | sMATF mirror | [11 empty] | - | 8 | 4.0239 |  | - | -0.00665 | 0.0084 |  |  | Average of initial/closeout
2023-06-14 | - | EB | BA7 | Y | sMATF LED pri | [11 empty] | - | [6 empty] | 814.167 | 597.8305 | Average of initial/closeout
2023-06-15 | 11:52 | MW | BA8 |  |  | 20.8735 | 23.811 | 21.8164 | 24.149 | 24.439 | 23.413 | [9 empty] | 15255 |  |  |  |  | Take dark image and enable dark subtract after starting up ABMC.vi
2023-06-15 | 12:53 | MW | BA8 |  | sMATF mirror | 20.8735 | 23.811 | 21.8164 | 24.149 | 24.439 | 23.413 | [5 empty] | 0,19 |  |  |  | 125103 | -0.0046 | 0.0028 |  |  | Leica beam centroid is ~ at 0,16 focus, standard deviation ~um
2023-06-15 |  | MW | BA8 |  | sMATF LED pri | 20.8735 | 23.811 | 21.8164 | 24.149 | 24.439 | 23.413 | [12 empty] | 809.382 | 589.557
2023-06-15 | 13:36 | MW | BA8 |  | sMATF mirror | 20.748 | 23.8192 | 21.8236 | 24.0398 | 24.432 | 23.42 | [5 empty] | 0,19 | 14 | 4.0241 |  | 34833 | 0.002 | -0.001 |  |  | Leica beam centroid is ~ 800/584 at 0,16 focus, standard deviation ~7um
2023-06-15 | 13:22 | MW | BA8 |  | sMATF LED pri | 20.748 | 23.8192 | 21.8236 | 24.0398 | 24.432 | 23.42 | [5 empty] | 1,79 |  |  |  | 34643 |  |  | 807.579 | 602.536
2023-06-15 | 14:20 | MW | BA8 |  | sMATF mirror | 20.748 | 23.8192 | 21.8236 | 24.0398 | 24.432 | 23.42 | [5 empty] | 0,19 |  |  |  | 42104 | -0.0035 | 0.0056
2023-06-15 | 14:25 | MW | BA8 |  | sMATF LED pri | 20.748 | 23.8192 | 21.8236 | 24.0398 | 24.432 | 23.42 | [5 empty] | 1,79 |  |  |  | 42352 |  |  | 813.565 | 598.427
2023-06-15 | 14:46 | MW | BA8 |  | sMATF LED pri | 20.748 | 23.8192 | 21.8236 | 24.0398 | 24.432 | 23.42 | [5 empty] | 1,79 |  |  |  | 44450 |  |  | 811.563 | 598.677
2023-06-15 | 15:30 | MW | BA8 |  | sMATF LED pri | 20.748 | 23.8192 | 21.8236 | 24.0398 | 24.432 | 23.42 | [5 empty] | 1,79 |  |  |  | 52831 |  |  | 813.385 | 596.834 | X centroid tool has a >2-pix variation, reason TBD
2023-06-15 | 15:49 | MW | BA8 |  | sMATF mirror | 20.748 | 23.8192 | 21.8236 | 24.0398 | 24.432 | 23.42 | [5 empty] | 0,19 |  |  |  | 54716 | -0.0042 | 0.006
2023-06-15 | 16:13 | MW | BA8 |  | sMATF LED pri | 20.748 | 23.8192 | 21.8236 | 24.0398 | 24.432 | 23.42 | [5 empty] | 1,79 |  |  |  | 61017 |  |  | 812.699 | 597.569 | no dark subtract
2023-06-15 | 16:18 | MW | BA8 |  | sMATF LED pri | 20.748 | 23.8192 | 21.8236 | 24.0398 | 24.432 | 23.42 | [5 empty] | 1,80 |  |  |  | 61634 |  |  | 813.319 | 596.859 | dark subtract
2023-06-15 |  | MW | BA8 |  | sMATF LED pri | 20.748 | 23.8192 | 21.8236 | 24.0398 | 24.432 | 23.42 | [5 empty] | 1,81 |  |  |  | 61844 |  |  | 813.358 | 596.826 | dark subtract
2023-06-15 |  | MW | BA8 |  | sMATF LED pri | 20.748 | 23.8192 | 21.8236 | 24.0398 | 24.432 | 23.42 | [5 empty] | 1,82 |  |  |  | 62122 |  |  | 811.624 | 596.943 | dark subtract
2023-06-15 | 16:32 | MW | BA8 |  | sMATF LED pri | 20.748 | 23.8192 | 21.8236 | 24.0398 | 24.432 | 23.42 | [5 empty] | 1,83 |  |  |  | 62953 |  |  | 812.681 | 597.662 | no dark subtract
2023-06-15 | 16:34 | MW | BA8 |  | sMATF mirror | 20.748 | 23.8192 | 21.8236 | 24.0398 | 24.432 | 23.42 | [5 empty] | 0,19 |  |  |  | 63204 | -0.0042 | 0.0056
2023-06-15 | 17:25 | MW | BA8 | Y | PATB mirror | 20.748 | 23.8192 | 21.8236 | 24.0398 | 24.441 | 23.413 | 101.518 | 125.739 | 0.001 | 0.1047 | 0.0822 | 0,19 | 13 | 3.4744 |  | 71743 | -0.0049 | 0.0056 |  |  | Std dev~6um
2023-06-15 | 17:50 | MW | BA8 | Y | PATB LED pri | 20.748 | 23.8192 | 21.8236 | 24.0398 | 24.441 | 23.413 | 88.796 | 108.686 | 0.002 | 0.1047 | 0.0822 | 1135 |  |  |  | 74704 |  |  | 812.201 | 598.029
2023-06-15 | 18:13 | MW | BA8 |  | sMATF mirror | 20.748 | 23.8192 | 21.8236 | 24.0398 | 24.431 | 23.411 | -2147 | -65 | 0 | 0 | 0 | 0,19 | 11 | 4.0241 |  | 80614 | -0.0042 | 0.0056 |  |  | Std dev 9um
2023-06-15 | 18:21 | MW | BA8 |  | sMATF LED pri | 20.748 | 23.8192 | 21.8236 | 24.0398 | 24.431 | 23.411 | -2147 | -65 | 0 | 0 | 0 | 1,79 |  |  |  | 81841 |  |  | 813.53 | 597.584
2023-06-15 |  | EB | BA8 | Y | sMATF mirror | [11 empty] | - | 12 | 4.0241 |  | - | -0.0042 | 0.0056 |  |  | Average of initial/closeout
2023-06-15 |  | EB | BA8 | Y | sMATF LED pri | [11 empty] | - | [6 empty] | 813.1055 | 597.623 | Average of initial/closeout
2023-06-16 | 17:54 | MW | NULLING |  | sMATF mirror | 20.8771 | 23.8137 | 21.8254 | 24.1543 | 24.428 | 23.411 | -2147.484 | -65.003 | -60.015 | -0.0001 | 0 | 0,21 | 13 | 4.0239 |  | 74328 | -0.0015 | 0.0025 |  |  | Std Dev ~9 um
2023-06-16 | 17:40 | MW | NULLING |  | sMATF LED pri | 20.8771 | 23.8137 | 21.8254 | 24.1543 | 24.428 | 23.411 | -2147.484 | -65.003 | -60.015 | -0.0001 | 0 | 1,76 |  |  |  | 73900 |  |  | 807.624 | 601.66
2023-06-16 | 18:16 | MW | SETTLING |  | sMATF mirror | 20.8771 | 23.8137 | 21.8254 | 24.1543 | 24.428 | 23.411 | -2147.484 | -65.003 | -60.015 | -0.0001 | 0 | 0,25 |  |  |  | 81454 | -0.0022 | 0.0035
2023-06-16 | 18:20 | MW | SETTLING |  | sMATF LED pri | 20.8771 | 23.8137 | 21.8254 | 24.1543 | 24.428 | 23.411 | -2147.484 | -65.003 | -60.015 | -0.0001 | 0 | 1,77 |  |  |  | 81822 |  |  | 809.634 | 599.71
2023-06-16 | 19:21 | MN | SETTLING |  | sMATF mirror | 20.8771 | 23.8137 | 21.8254 | 24.1543 | 24.428 | 23.411 | -2147.484 | -65.003 | -60.015 | -0.0001 | 0 | 0,22 |  |  |  | 91626 | -0.0025 | 0.0035
2023-06-16 | 19:27 | MW | SETTLING |  | sMATF LED pri | 20.8771 | 23.8137 | 21.8254 | 24.1543 | 24.428 | 23.411 | -2147.484 | -65.003 | -60.015 | -0.0001 | 0 | 1,70 |  |  |  | 92458 |  |  | 808.6 | 601.533
2023-06-16 | 19:50 | MW | BA9 |  | sMATF LED pri | 20.8771 | 23.8137 | 21.8254 | 24.1543 | 24.428 | 23.411 | -2147.484 | -65.003 | -60.015 | -0.0001 | 0 | 1,70 |  |  |  | 94841 |  |  | 809.546 | 599.682
2023-06-16 | 20:00 | MW | BA9 |  | sMATF mirror | 20.8771 | 23.8137 | 21.8254 | 24.1543 | 24.428 | 23.411 | -2147.484 | -65.003 | -60.015 | -0.0001 | 0 | 0,22 | 13 | 4.024 |  | 95814 | -0.0035 | 0.0035 |  |  | Std Dev ~5 um
2023-06-16 | 20:28 | MW | BA9 | Y | PATB mirror | 20.8771 | 23.8137 | 21.8254 | 24.1543 | 24.428 | 23.411 | 101.521 | 125.742 | 0.001 | 0.10754 | 0.0882 | 0,22 | 23 | 3.4744 |  | 202544 | -0.0032 | 0.0035 |  |  | Std Dev ~9 um
2023-06-16 | 20:42 | MW | BA9 | Y | PATB LED pri | 20.8771 | 23.8137 | 21.8254 | 24.1543 | 24.428 | 23.411 | 88.729 | 108.209 | -0.003 | 0.1078 | 0.0884 | 1128 |  |  |  | 204432 |  |  | 809.766 | 599.7
2023-06-16 | 21:17 | MW | BA9 |  | sMATF mirror | 20.8771 | 23.8137 | 21.8254 | 24.1543 | 24.428 | 23.411 | -2147.484 | -65.003 | -60.015 | -0.0001 | 0 | 0,20 | 20 | 4.024 |  | 11439 | -0.0035 | 0.0025 |  |  | 6um
2023-06-16 | 21:20 | MW | BA9 |  | sMATF LED pri | 20.8771 | 23.8137 | 21.8254 | 24.1543 | 24.428 | 23.411 | -2147.484 | -65.003 | -60.015 | -0.0001 | 0 | 1,78 |  |  |  | 11843 |  |  | 808.708 | 599.639
2023-06-16 | - | EB | BA9 | Y | sMATF mirror | [11 empty] | - | - | 4.024 |  | - | -0.0035 | 0.003 |  |  | Average of initial/closeout
2023-06-16 | - | EB | BA9 | Y | sMATF LED pri | [11 empty] | - | [6 empty] | 809.127 | 599.6605 | Average of initial/closeout
2023-06-20 | 20:17 | MW | BA9 |  | sMATF LED pri | 20.8771 | 23.8137 | 21.8254 | 24.1543 | 24.428 | 23.411 | -2147.484 | -65.003 | -60.015 | -0.0001 | 0 | 1,70 |  |  |  | 201701 |  |  | 830.573 | 605.601
2023-06-20 | 20:40 | MW | BA9 |  | sMATF mirror | 20.8771 | 23.8137 | 21.8254 | 24.1543 | 24.428 | 23.411 | -2147.484 | -65.003 | -60.015 | -0.0001 | 0 | 0,15 |  |  |  | 203946 | 0.0034 | 0.0313
2023-06-20 | 21:39 | MW | BA9 |  | sMATF LED pri | 21.3771 | 23.8137 | 21.8254 | 24.2043 | [7 empty] | 1,70 |  |  |  | 13731 |  |  | 836.531 | 10.5201
2023-06-20 |  | MW | BA9 |  | sMATF LED pri | 20.3771 | 23.8137 | 21.8254 | 24.1043 | [7 empty] | 1,70 |  |  |  | 14229 |  |  | 842.449 | 1192.01
2023-06-20 |  | MW | BA9 |  | sMATF LED pri | 20.8771 | 24.3137 | 21.8254 | 24.1543 | [7 empty] | 1,70 |  |  |  | 14290 |  |  | 1586.7 | 595.6
2023-06-20 | 22:05 | MW | BA9 |  | sMATF LED pri | 20.8771 | 23.3137 | 21.8254 | 24.1543 | [7 empty] | 1,70 |  |  |  | 20719 |  |  | 83.66 | 605.65
2023-06-20 |  | MW | BA9 |  | sMATF LED pri | 20.8771 | 23.8137 | 22.5254 | 24.1543 | [7 empty] | 1,70 |  |  |  | 21226 |  |  | 108.6 | 623.8
2023-06-20 | 22:18 | MW | BA9 |  | sMATF LED pri | 20.8771 | 23.8137 | 21.1254 | 24.1543 | [7 empty] | 1,70 |  |  |  | 21605 |  |  | 1549.6 | 600.5
2023-06-20 | 22:23 | MW | BA9 |  | sMATF LED pri | 20.9771 | 23.8137 | 21.8254 | 24.6543 | [7 empty] | 1,70 |  |  |  | 22241 |  |  | 846.6 | 1189.4
2023-06-20 | 22:28 | MW | BA9 |  | sMATF LED pri | 20.7771 | 23.8137 | 21.8254 | 23.654 | [7 empty] | 1,70 |  |  |  | 22628 |  |  | 821.6 | 4.5
2023-06-20 | 22:40 | MW | BA9 |  | sMATF mirror | 21.3771 | 23.8137 | 21.8254 | 24.55 | [7 empty] | 0,15 |  |  |  | 23811 | -0.0133 | 0.0424
2023-06-20 | 22:50 | MW | BA9 |  | sMATF mirror | 20.3771 | 23.8137 | 21.8254 | 23.7043 | [7 empty] | 0,15 |  |  |  | 24829 | -0.0226 | 0.0195
2023-06-20 | 22:56 | MW | BA9 |  | sMATF mirror | 20.8771 | 24.3137 | 22.5754 | 24.154 | [7 empty] | 0,15 |  |  |  | 25355 | 0.021 | -0.0395
2023-06-20 | 23:03 | MW | BA9 |  | sMATF mirror | 20.8771 | 23.3137 | 21.1254 | 24.1543 | [7 empty] | 0,15 |  |  |  | 30124 | -0.005 | 0.041
2023-06-20 | 23:08 | MW | BA9 |  | sMATF mirror | 20.8771 | 24.2637 | 22.5254 | 24.1543 | [7 empty] | 0,15 |  |  |  | 30559 | 0.014 | -0.042
2023-06-20 | 23:16 | MW | BA9 |  | sMATF mirror | 20.8771 | 23.3137 | 21.1254 | 24.1543 | [7 empty] | 0,15 |  |  |  | 31049 | -0.006 | 0.06
2023-06-20 | 23:17 | MW | BA9 |  | sMATF mirror | 21.4271 | 23.8137 | 21.8254 | 24.6543 | [7 empty] | 0,15 |  |  |  | 31526 | 0.031 | 0.046
2023-06-20 | 23:25 | MW | BA9 |  | sMATF mirror | 20.3271 | 23.8137 | 21.8254 | 23.654 | [7 empty] | 0,15 |  |  |  | 32252 | -0.0654 | 0.0216
2023-06-20 | 23:26 | MW | BA9 |  | sMATF mirror | 20.8771 | 23.8137 | 21.8254 | 24.1543 | 24.428 | 23.411 | -2147.484 | -65.003 | -60.015 | -0.0001 | 0 | 0,15 |  |  |  | 32531 | -0.0001 | 0.0317
2023-06-20 | 23:27 | MW | BA9 |  | sMATF LED pri | 20.8771 | 23.8137 | 21.8254 | 24.1543 | 24.428 | 23.411 | -2147.484 | -65.003 | -60.015 | -0.0001 | 0 | 1,70 |  |  |  | 32739 |  |  | 832.99 | 603.669
[12 empty] | -2147.484 | 0.4 | -2147.484 | 0 | 0
2023-07-21 | 22:39 | CRD | post ADM leak test |  | sMATF retro | 20.3323 | 23.6257 | 21.5334 | 23.6674 | 24.153 | 23.75 | [6 empty] | 23 | 4.0341 |  |  | NA | NA |  |  | just a functional check to prove that things are working well after PR3509-018, don't care about making measurements in tolerance. Doing it in STOW
2023-07-21 | 22:24 | CRD | post ADM leak test |  | sMATF mirror | 20.8771 | 23.8137 | 21.8254 | 24.1543 | 24.428 | 23.411 | -2147.484 | 0.4 | -2147.484 | 0 | 0 | 0,20 | 1 | NA |  | 042426 | 0.0003 | -0.0152
2023-07-21 | 22:35 | CRD | post ADM leak test |  | sMATF LED pri | 20.8771 | 23.8137 | 21.8254 | 24.1543 | 24.428 | 23.411 | -2147.484 | 0.4 | -2147.484 | 0 | 0 | 1,70 |  |  |  | 043538 |  |  | 959.7 | 609.3
preTVAC0 warm vacuum (WOA 4014 event 70)
2023-07-24 | 6:56 | MP+CRD | nulling |  | sMATF mirror | 20.9241 | 23.7769 | 21.8032 | 24.1839 | 24.439 | 23.42 | [5 empty] | 0,20 | 23 | 4.0232 |  | 065629 | -0.0001 | 0.0001
2023-07-24 | 6:54 | MP+CRD | nulling |  | sMATF LED pri | 20.9241 | 23.7769 | 21.8032 | 24.1839 | 24.439 | 23.42 | [5 empty] | 1,76 |  |  |  | 065403 |  |  | 807.89 | 602.49
2023-07-24 | 7:14 | MP | Ti1C2 | Y | sMATF retro | 20.9241 | 23.7769 | 21.8032 | 24.1839 | 24.153 | 23.75 | [6 empty] | 42 | 4.0338 | [6 empty] | 23 C, 0 mbar
2023-07-24 | 7:16 | MP | Ti1C2 | Y | CHIPS retro | 20.9241 | 23.7769 | 21.8032 | 24.1839 | 24.63 | 23.18 | [6 empty] | 23 | 4.7646 | [6 empty] | 23 C, 0 mbar
2023-07-24 | 7:20 | MP | settling |  | sMATF mirror | 20.9241 | 23.7769 | 21.8032 | 24.1839 | 24.439 | 23.42 | [5 empty] | 0,20 |  |  |  | 072049 | -0.0022 | 0.0021
2023-07-24 | 7:22 | MP | settling |  | sMATF LED pri | 20.9241 | 23.7769 | 21.8032 | 24.1839 | 24.439 | 23.42 | [5 empty] | 1,76 |  |  |  | 072252 |  |  | 809.97 | 600.3
2023-07-24 | 7:43 | MP | settling |  | sMATF mirror | 20.9241 | 23.7769 | 21.8032 | 24.1839 | 24.439 | 23.42 | [5 empty] | 0,20 |  |  |  | 074329 | -0.0025 | 0.0025 | 0 | 0
2023-07-24 | 7:41 | MP | settling |  | sMATF LED pri | 20.9241 | 23.7769 | 21.8032 | 24.1839 | 24.439 | 23.42 | [5 empty] | 1,76 |  |  |  | 074120 |  |  | 810.44 | 599.89 |  | Background Subtraction checkout |  |  |  | 080041 image
2023-07-24 | 7:58 | MP | settling |  | sMATF mirror | 20.9241 | 23.7769 | 21.8032 | 24.1839 | 24.439 | 23.42 | [5 empty] | 0,20 |  |  |  | 075848 | -0.0028 | 0.0025 | 0.47 | -0.41 |  | Crosshair X | Crosshair Y | Centroid X | Centroid Y
2023-07-24 | 8:00 | MP | settling |  | sMATF LED pri | 20.9241 | 23.7769 | 21.8032 | 24.1839 | 24.439 | 23.42 | [5 empty] | 1,76 |  |  |  | 080041 |  |  | 810.55 | 599.83 | 810600 | 810 | 600 | 810.55 | 599.83
2023-07-24 | 8:26 | MP | Ti1 |  | sMATF mirror | 20.9241 | 23.7769 | 21.8032 | 24.1839 | 24.439 | 23.42 | -2147.484 | -65.003 | -60.015 | -0.0001 | 0 | 0,20 | 46 | 4.0232 |  | 082608 | -0.0025 | 0.0025 | 0.58 | -0.47 |  | 806 | 599 | 810.41 | 599.74
2023-07-24 | 8:23 | MP | Ti1 |  | sMATF LED pri | 20.9241 | 23.7769 | 21.8032 | 24.1839 | 24.439 | 23.42 | -2147.484 | -65.003 | -60.015 | -0.0001 | 0 | 1,76 |  |  |  | 082336 |  |  | 810.69 | 599.68 |  | 815 | 600 | 810.62 | 599.82
2023-07-24 | 8:59 | MP | Ti1 | Y | PATB mirror | 20.9241 | 23.7769 | 21.8032 | 24.1839 | 24.439 | 23.42 | 101.52 | 125.742 | -0.003 | 0.103 | 0.0875 | 0,20 | 75 | 3.474 |  | 085956 | -0.0025 | 0.0025 | 0.72 | -0.62 |  | 809 | 600 | 810.55 | 599.814
2023-07-24 | 9:24 | MP | Ti1 | Y | PATB LED pri | 20.9241 | 23.7769 | 21.8032 | 24.1839 | 24.439 | 23.42 | 88.613 | 108.247 | -0.002 | 0.1031 | 0.0875 | 1132 |  |  |  | 092435 |  |  | 810.95 | 599.76 |  | 808 | 600 | 810.535 | 599.802
2023-07-24 | 9:34 | MP | Ti1 | Y | PATB LED red | 20.9241 | 23.7769 | 21.8032 | 24.1839 | 24.439 | 23.42 | 111.598 | 148.726 | 0 | 0.1028 | 0.0875 | 1132 |  |  |  | 093441 |  |  | 810.69 | 599.43 |  | 807 | 600 | 810.503 | 599.781
2023-07-24 | 9:47 | MP | Ti1 | Y | PATB retro pri | 20.9241 | 23.7769 | 21.8032 | 24.1839 | 24.439 | 23.42 | 111.598 | 148.608 | 0 | 0.1028 | 0.0875 | 0,20 | 45 | 3.4605 | [6 empty] | 23 C, 0 mbar | 811 | 600 | 810.549 | 599.836
2023-07-24 |  | MP | Ti1 |  | PATB retro red | 20.9241 | 23.7769 | 21.8032 | 24.1839 | 24.439 | 23.42 | 92.911 | 144.902 | 0 | 0.1027 | 0.0876 | 0,20 | 1 | No Return, pose doesn't compare wrt previous meas. | [6 empty] | 23 C, 0 mbar | 812 | 600 | 810.557 | 599.835
2023-07-24 | 10:28 | MP | Ti1 | Y | PATA mirror | 20.9241 | 23.7769 | 21.8032 | 24.1839 | 24.439 | 23.42 | -76.119 | 125.702 | 0 | 0.0797 | 0.077 | 0,20 | 35 | 3.474 |  | 102807 | -0.0025 | 0.0028 |  |  |  | 813 | 600 | 810.552 | 599.84
2023-07-24 | 10:39 | MP | Ti1 | Y | PATA LED pri | 20.9241 | 23.7769 | 21.8032 | 24.1839 | 24.439 | 23.42 | -60.476 | 108.224 | 0 | 0.0797 | 0.077 | 1128 |  |  |  | 103924 |  |  | 810.78 | 599.86 |  | 814 | 600 | 810.602 | 599.827
2023-07-24 | 10:56 | MP | Ti1 | Y | PATA LED red | 20.9241 | 23.7769 | 21.8032 | 24.1839 | 24.439 | 23.42 | -83.552 | 148.482 | 0 | 0.0794 | 0.077 | 1128 |  |  |  | 105655 |  |  | 810.57 | 599.3
2023-07-24 | 11:04 | MP | Ti1 | Y | PATA retro pri | 20.9241 | 23.7769 | 21.8032 | 24.1839 | 24.439 | 23.42 | -72.08 | 98.979 | 0 | 0.0799 | 0.077 | 0,20 | 19 | 3.4604 | [6 empty] | 23 C, 0 mbar
2023-07-24 |  | MP | Ti1 |  | PATA retro red | 20.9241 | 23.7769 | 21.8032 | 24.1839 | 24.439 | 23.42 | -26.991 | 144.602 | 0 | 0.0795 | 0.077 | 0,20 | 1 | No Return | [6 empty] | 23 C, 0 mbar
2023-07-24 | 11:31 | MP | Ti1 |  | sMATF mirror | 20.9241 | 23.7769 | 21.8032 | 24.1839 | 24.439 | 23.42 | -2147.484 | -65.003 | -60.015 | -0.0001 | 0 | 0,20 | 15 | 4.0233 |  | 113146 | -0.0032 | 0.0028
2023-07-24 | 11:35 | MP | Ti1 |  | sMATF LED pri | 20.9241 | 23.7769 | 21.8032 | 24.1839 | 24.439 | 23.42 | -2147.484 | -65.003 | -60.015 | -0.0001 | 0 | 1,76 |  |  |  | 113537 |  |  | 811.02 | 599.47
2023-07-24 | 12:12 | MP | STOW cross check |  | sMATF mirror | 20.9241 | 23.7769 | 21.8032 | 24.1839 | 24.439 | 23.42 | -2147.484 | -64 | -2147.484 | -0.0001 | 0.0002 | 0,20 | 15 | 4.0233 |  | 121242 | -0.0025 | 0.0028 | 0.33 | -0.21 | These correspond with WOA4014 event 80, STOW vs H2 cross check
2023-07-24 | 12:16 | MP | STOW cross check |  | sMATF LED pri | 20.9241 | 23.7769 | 21.8032 | 24.1839 | 24.439 | 23.42 | -2147.484 | -64 | -2147.484 | -0.0001 | 0.0002 | 1,76 |  |  |  | 121637 |  |  | 810.76 | 599.46
2023-07-24 | 12:46 | MP | ADM cross check |  | sMATF mirror | 20.9241 | 23.7769 | 21.8032 | 24.1839 | 24.439 | 23.42 | -2147.484 | -65.003 | -60.015 | -0.0001 | 0 | 0,20 | 19 | 4.0233 |  | 124657 | -0.0032 | 0.0028
2023-07-24 | 12:51 | MP | ADM cross check |  | sMATF LED pri | 20.9241 | 23.7769 | 21.8032 | 24.1839 | 24.439 | 23.42 | -2147.484 | -65.003 | -60.015 | -0.0001 | 0 | 1,76 |  |  |  | 125156 |  |  | 811.02 | 599.51
|  | EB | Ti1 | Y | sMATF mirror | [13 empty] | 4.0233 |  |  | -0.00285 | 0.00265 | 0.53 | -0.12 | Average of sMATF measurments
|  | EB | Ti1 | Y | sMATF LED pri | [18 empty] | 810.855 | 599.575
2023-08-02 | 15:23 | MP | INITIAL |  | sMATF mirror | 0 | 0 | 0 | 0 | 0.01 | 0.005 | -2147.484 | -65.003 | -60.015 | -0.0001 | 0 | 0,20 | 15 | 4.0235 |  | 212303 | 0.0163 | 0.016 |  |  | 639.4 torr (852.5 mbar)
2023-08-02 | 15:27 | MP | INITIAL |  | sMATF LED pri | 0 | 0 | 0 | 0 | 0.01 | 0.005 | -2147.484 | -65.003 | -60.015 | -0.0001 | 0 | 1,76 |  |  |  | 212704 |  |  | 819.5 | 612.1
2023-08-02 | 15:23 | MP | after Homed |  | sMATF mirror | 20.9241 | 23.7769 | 21.8032 | 24.1839 | 24.439 | 23.42 | -2147.484 | -65.003 | -60.015 | -0.0001 | 0 | 0,20 |  |  |  | 212303 | 0.0076 | -0.0235
2023-08-02 | 16:02 | MP | after Homed |  | sMATF LED pri | 20.9241 | 23.7769 | 21.8032 | 24.1839 | 24.439 | 23.42 | -2147.484 | -65.003 | -60.015 | -0.0001 | 0 | 1,76 |  |  |  | 220216 |  |  | 785.3 | 604.8
2023-08-02 | 17:02 | MP | Nulled |  | sMATF mirror | 20.8752 | 23.8404 | 21.8697 | 24.133 | 24.439 | 23.42 | -2147.484 | -65.003 | -60.015 | -0.0001 | 0 | 0,20 |  |  |  | 230248 | 0.0013 | -0.0027
2023-08-02 | 17:07 | MP | Nulled |  | sMATF LED pri | 20.8752 | 23.8404 | 21.8697 | 24.133 | 24.439 | 23.42 | -2147.484 | -65.003 | -60.015 | -0.0001 | 0 | 1,76 |  |  |  | 230741 |  |  | 807.9 | 602.71
2023-08-02 | 23:21 | MP,ML | Ti2 | Y | sMATF retro | 20.8752 | 23.8404 | 21.8697 | 24.133 | 24.153 | 23.75 | -2147.484 | -65.003 | -60.015 | -0.0001 | 0 |  | 24 | 4.0339
2023-08-02 | 23:32 | MP,ML | Ti2 | Y | CHIPS retro | 20.8752 | 23.8404 | 21.8697 | 24.133 | 24.63 | 23.18 | -2147.484 | -65.003 | -60.015 | -0.0001 | 0 |  | 15 | 4.7648
2023-08-02 | 23:37 | MP,ML | Settling |  | sMATF mirror | 20.8752 | 23.8404 | 21.8697 | 24.133 | 24.432 | 23.41 | -2147.484 | -65.003 | -60.015 | -0.0001 | 0 | 0,20 | 13 | 4.0234 |  | 233700 | -0.0022 | 0.0011
2023-08-02 | 23:36 | MP,ML | Settling |  | sMATF LED pri | 20.8752 | 23.8404 | 21.8697 | 24.133 | 24.432 | 23.41 | -2147.484 | -65.003 | -60.015 | -0.0001 | 0 | 1,76 |  |  |  | 233601 | 0 | 0 | 810.18 | 600.24
2023-08-02 | 23:52 | MP,ML | Settling |  | sMATF mirror | 20.8752 | 23.8404 | 21.8697 | 24.133 | 24.432 | 23.41 | -2147.484 | -65.003 | -60.015 | -0.0001 | 0 | 0,20 |  |  |  | 235244 | -0.0022 | 0.0011 | 0 | 0
2023-08-02 | 23:53 | MP,ML | Settling |  | sMATF LED pri | 20.8752 | 23.8404 | 21.8697 | 24.133 | 24.432 | 23.41 | -2147.484 | -65.003 | -60.015 | -0.0001 | 0 | 1,76 |  |  |  | 235352 | 0 | 0 | 810.26 | 600.34
2023-08-03 | 0:08 | MP,ML | Settling |  | sMATF mirror | 20.8752 | 23.8404 | 21.8697 | 24.133 | 24.432 | 23.41 | -2147.484 | -65.003 | -60.015 | -0.0001 | 0 | 0,20 |  |  |  | 000823 | -0.0025 | 0.0011 | 0.08 | 0.1
2023-08-03 | 0:06 | MP,ML | Settling |  | sMATF LED pri | 20.8752 | 23.8404 | 21.8697 | 24.133 | 24.432 | 23.41 | -2147.484 | -65.003 | -60.015 | -0.0001 | 0 | 1,76 |  |  |  | 000632 | -0.0003 | 0 | 810.22 | 599.9
2023-08-03 | 0:22 | MP,ML | Settling |  | sMATF mirror | 20.8752 | 23.8404 | 21.8697 | 24.133 | 24.432 | 23.41 | -2147.484 | -65.003 | -60.015 | -0.0001 | 0 | 0,20 |  |  |  | 002246 | -0.0028 | 0.0011 | 0.04 | -0.34
2023-08-03 | 0:23 | MP,ML | Settling |  | sMATF LED pri | 20.8752 | 23.8404 | 21.8697 | 24.133 | 24.432 | 23.41 | -2147.484 | -65.003 | -60.015 | -0.0001 | 0 | 1,76 |  |  |  | 002339 | -0.0006 | 0 | 810.35 | 599.71
2023-08-03 | 0:39 | MP,ML | Ti2 |  | sMATF mirror | 20.8752 | 23.8404 | 21.8697 | 24.133 | 24.432 | 23.41 | -2147.484 | -65.003 | -60.015 | -0.0001 | 0 | 0,20 | 20 | 4.0234 |  | 003946 | -0.0025 | 0.0018 | 0.17 | -0.53
2023-08-03 | 0:37 | MP,ML | Ti2 |  | sMATF LED pri | 20.8752 | 23.8404 | 21.8697 | 24.133 | 24.432 | 23.41 | -2147.484 | -65.003 | -60.015 | -0.0001 | 0 | 1,76 |  |  |  | 003739 | -0.0003 | 0.0007 | 810.49 | 599.63
2023-08-03 | 1:01 | MP,ML | prenull |  | PATB mirror | 20.8752 | 23.8404 | 21.8697 | 24.133 | 24.432 | 23.41 | 101.52 | 125.742 | -0.003 | 0.1028 | 0.0878 | 0,20 |  |  |  | 010106 | -0.0049 | -0.0037
2023-08-03 | 1:08 | MP,ML | Ti2 | Y | PATB mirror | 20.8752 | 23.8404 | 21.8697 | 24.133 | 24.432 | 23.41 | 101.52 | 125.742 | -0.003 | 0.1077 | 0.0905 | 0,20 | 40 | 3.474 |  | 010844 | -0.0028 | 0.0021
2023-08-03 | 1:17 | MP,ML | prenull |  | PATB LED pri | 20.8752 | 23.8404 | 21.8697 | 24.133 | 24.432 | 23.41 | 88.612 | 108.247 | -0.001 | 0.1077 | 0.0905 | 1132 |  |  |  | 011728 |  |  | 811.27 | 599.87
2023-08-03 | 1:22 | MP,ML | Ti2 | Y | PATB LED pri | 20.8752 | 23.8404 | 21.8697 | 24.133 | 24.432 | 23.41 | 88.633 | 108.217 | -0.001 | 0.1077 | 0.0905 | 1132 |  |  |  | 012248 |  |  | 811.86 | 600.2
2023-08-03 | 1:33 | MP,ML | prenull |  | PATA mirror | 20.8752 | 23.8404 | 21.8697 | 24.133 | 24.432 | 23.41 | -76.12 | 125.703 | -0.002 | 0.0795 | 0.0767 | 0,20 |  |  |  | 013351 | -0.006 | 0.0008
2023-08-03 | 1:41 | MP,ML | Ti2 | Y | PATA mirror | 20.8752 | 23.8404 | 21.8697 | 24.133 | 24.432 | 23.41 | -76.12 | 125.703 | -0.002 | 0.0826 | 0.0755 | 0,20 | 35 | 3.4739 |  | 014124 | -0.0028 | 0.0015
2023-08-03 | 1:51 | MP,ML | prenull |  | PATA LED pri | 20.8752 | 23.8404 | 21.8697 | 24.133 | 24.432 | 23.41 | -60.476 | 108.224 | 0 | 0.0826 | 0.0755 | 1131 |  |  |  | 015141 |  |  | 810.85 | 600.62
2023-08-03 | 2:03 | MP,ML | Ti2 | Y | PATA LED pri | 20.8752 | 23.8404 | 21.8697 | 24.133 | 24.432 | 23.41 | -60.476 | 108.224 | 0 | 0.0826 | 0.0755 | 1131 |  |  |  | 020353 |  |  | 810.45 | 599.86
2023-08-03 | 2:21 | MP,ML | Ti2 |  | sMATF mirror | 20.8752 | 23.8404 | 21.8697 | 24.133 | 24.432 | 23.41 | -2147.484 | -64.9 | -60.012 | -0.0003 | 0 | 0,20 | 25 | 4.0234 |  | 022102 | -0.0028 | 0.0018
2023-08-03 | 2:23 | MP,ML | Ti2 |  | sMATF LED pri | 20.8752 | 23.8404 | 21.8697 | 24.133 | 24.432 | 23.41 | -2147.484 | -64.9 | -60.012 | -0.0003 | 0 | 1,76 |  |  |  | 022340 | -0.0003 | 0 | 810.45 | 599.24
2023-08-03 |  | EB | Ti2 | Y | sMATF mirror | [13 empty] | 4.0234 |  |  | -0.00265 | 0.0018 | -0.04 | -0.39 | Average of sMATF measurments
2023-08-03 |  | EB | Ti2 | Y | sMATF LED pri | [18 empty] | 810.47 | 599.435
2023-08-06 | 16:05 | MP,ML | INITIAL |  | sMATF mirror | 0 | 0 | 0 | 0 | 0 | 0 | [5 empty] | 0,20 | 15 | 4.0168 |  | 160529 | -0.0049 | -0.0027 |  |  | 16:28 rebooted old 5DOF computer to reconnect to Lecia
2023-08-06 | 16:17 | MP,ML | INITIAL |  | sMATF LED pri | 0 | 0 | 0 | 0 | 0 | 0 | [5 empty] | 1,76 |  | -0.006579 |  | 161727 |  |  | 782.4 | 564.9 | ADM int1 temp=295, press=855 mbar(641.3 torr)
[6 empty] | REBOOT RM CONTROL | [19 empty] | REBOOTED new 5DOF, power cycled RM & ethernet controllers
2023-08-06 | 18:28 | MP,ML | after Homed |  | sMATF mirror | 20.8752 | 23.8404 | 21.8697 | 24.133 | 24.432 | 23.41 | [5 empty] | 0,20 |  |  |  | 182821 | 0.0027 | -0.0069
2023-08-06 | 18:32 | MP,ML | after Homed |  | sMATF LED pri | 20.8752 | 23.8404 | 21.8697 | 24.133 | 24.432 | 23.41 | [5 empty] | 1,76 |  |  |  | 183237 |  |  | 778.7 | 575
2023-08-06 | 19:49 | MP,ML | Nulled |  | sMATF mirror | 20.46 | 24.1454 | 22.2773 | 23.7641 | 24.436 | 23.414 | [5 empty] | 0,20 | 14 | 4.0167 |  | 194956 | -0.0008 | 0.0008
2023-08-06 | 19:51 | MP,ML | Nulled |  | sMATF LED pri | 20.46 | 24.1454 | 22.2773 | 23.7641 | 24.436 | 23.414 | [5 empty] | 1,76 |  |  |  | 195106 |  |  | 808.55 | 601.64
2023-08-06 | 19:57 | MP,ML | Ti3 | Y | sMATF retro | 20.46 | 24.1454 | 22.2773 | 23.7641 | 24.153 | 23.75 | [5 empty] | 0,20 | 13 | 4.0272 | [6 empty] | Average distance delta is -6.6 mm from warm measure
2023-08-06 | 19:59 | MP,ML | Ti3 | Y | CHIPS retro | 20.46 | 24.1454 | 22.2773 | 23.7641 | 24.63 | 23.18 | [5 empty] | 0,20 | 20 | 4.7582
2023-08-06 | 20:16 | MP,ML | Settling |  | sMATF mirror | 20.4588 | 24.1438 | 22.2773 | 23.7641 | 24.436 | 23.414 | [5 empty] | 0,20 |  |  |  | 201612 | 0.0003 | -0.0006 |  |  | ** Adjusting RM1 brought align back in like on 8/2 | -0.0042 | 0.0042
2023-08-06 | 20:14 | MP,ML | Settling |  | sMATF LED pri | 20.4588 | 24.1438 | 22.2773 | 23.7641 | 24.436 | 23.414 | [5 empty] | 1,76 |  |  |  | 201453 | 0.0025 | -0.0028 | 807.76 | 602.8 |  | 0 | 0.0031 | 811.05 | 599.31
2023-08-06 | 20:40 | MP,ML | Settling |  | sMATF mirror | 20.4588 | 24.1438 | 22.2773 | 23.7641 | 24.436 | 23.414 | [5 empty] | 0,20 |  |  |  | 204054 | -0.0022 | 0.0022 | -1.21 | 2.005
2023-08-06 | 20:41 | MP,ML | Settling |  | sMATF LED pri | 20.4588 | 24.1438 | 22.2773 | 23.7641 | 24.436 | 23.414 | [5 empty] | 1,76 |  |  |  | 204150 | 0 | 0 | 808.97 | 600.795 |  | Nulling.. | Az (deg) | EL (deg) | Px (pix) | Py (pix)
2023-08-06 | 21:01 | MP,ML | Settling |  | sMATF mirror | 20.4588 | 24.1438 | 22.2773 | 23.7641 | 24.436 | 23.414 | [5 empty] | 0,20 |  |  |  | 210151 | -0.0025 | 0.0025 | 0 | 0 |  | sMATF | -0.0008 | 0.0008 | -0.33 | -0.44
2023-08-06 | 21:03 | MP,ML | Settling |  | sMATF LED pri | 20.4588 | 24.1438 | 22.2773 | 23.7641 | 24.436 | 23.414 | [5 empty] | 1,76 |  |  |  | 210303 | -0.0003 | 0.0003 | 809.11 | 600.86 |  | PATA | 0 | 0 | -0.257 | 0.132
2023-08-06 | 21:21 | MP,ML | Settling |  | sMATF mirror | 20.4588 | 24.1438 | 22.2773 | 23.7641 | 24.436 | 23.414 | [5 empty] | 0,20 |  |  |  | 212134 | -0.0022 | 0.0025 | 0.14 | 0.065 |  | PATB | 0.0007 | -0.0003 | -0.279 | 0.067
2023-08-06 | 21:19 | MP,ML | Settling |  | sMATF LED pri | 20.4588 | 24.1438 | 22.2773 | 23.7641 | 24.436 | 23.414 | [5 empty] | 1,76 |  |  |  | 211948 | 0 | 0.0003 | 809.09 | 600.813
2023-08-06 | 21:39 | MP | Ti3 |  | sMATF mirror | 20.4588 | 24.1438 | 22.2773 | 23.7641 | 24.431 | 23.417 | [5 empty] | 0,20 | 19 | 4.0167 |  | 213947 | -0.0022 | 0.0025 | 0.12 | 0.018 | Adjusted RM3 and got better stdev.
2023-08-06 | 21:41 | MP | Ti3 |  | sMATF LED pri | 20.4588 | 24.1438 | 22.2773 | 23.7641 | 24.431 | 23.417 | [5 empty] | 1,76 |  | -3.5e-05 |  | 214124 | 0 | 0.0003 | 809.279 | 601.103 | ADM int1 temp=294.995, press=855 mbar(641.3 torr)
2023-08-06 | 22:12 | MP | prenull |  | PATB mirror | 20.4588 | 24.1438 | 22.2773 | 23.7641 | 24.431 | 23.417 | 101.522 | 125.748 | 0.002 | 0.1077 | 0.0906 | 0,20 |  |  |  | 221243 | 0.0208 | 0.0202 | 0.309 | 0.308
2023-08-06 | 22:20 | MP | Ti3 | Y | PATB mirror | 20.4588 | 24.1438 | 22.2773 | 23.7641 | 24.431 | 23.417 | 101.522 | 125.748 | 0.002 | 0.0788 | 0.09 | 0,20 | 29 | 3.4693 |  | 222054 | -0.0022 | 0.0025
2023-08-06 | 22:29 | MP | prenull |  | PATB LED pri | 20.4588 | 24.1438 | 22.2773 | 23.7641 | 24.431 | 23.417 | 88.629 | 108.216 | 0.002 | 0.079 | 0.09 | 1132 |  |  |  | 222904 |  |  | 824.882 | 617.613 | (spot was slightly off focus. 1,135 probably better)
2023-08-06 | 22:46 | MP | Ti3 | Y | PATB LED pri | 20.4588 | 24.1438 | 22.2773 | 23.7641 | 24.431 | 23.417 | 88.486 | 109.307 | 0.002 | 0.079 | 0.09 | 1132 |  |  |  | 224647 |  |  | 809.022 | 601.235 | (Camera reported error; restarted Labview)
2023-08-06 | 22:54 | MP | prenull |  | PATA mirror | 20.4588 | 24.1438 | 22.2773 | 23.7641 | 24.431 | 23.417 | -76.125 | 125.703 | -0.002 | 0.0823 | 0.0752 | 0,20 |  |  |  | 225431 | 0.0156 | -0.011
2023-08-06 | 23:01 | MP | Ti3 | Y | PATA mirror | 20.4588 | 24.1438 | 22.2773 | 23.7641 | 24.431 | 23.417 | -76.125 | 125.703 | -0.002 | 0.0517 | 0.0624 | 0,20 | 35 | 3.4693 |  | 230148 | -0.0015 | 0.0022 |  |  | (Whoops, measured range after LED)
2023-08-06 | 23:07 | MP | prenull |  | PATA LED pri | 20.4588 | 24.1438 | 22.2773 | 23.7641 | 24.431 | 23.417 | -60.474 | 108.266 | 0 | 0.0518 | 0.0625 | 1134 |  |  |  | 230751 |  |  | 824.742 | 616.065
2023-08-06 | 23:13 | MP | Ti3 | Y | PATA LED pri | 20.4588 | 24.1438 | 22.2773 | 23.7641 | 24.431 | 23.417 | -60.651 | 109.304 | 0 | 0.0518 | 0.0624 | 1134 |  |  |  | 231329 |  |  | 809 | 601.17
2023-08-06 | 23:37 | MP | Ti3 |  | sMATF mirror | 20.4588 | 24.1438 | 22.2773 | 23.7641 | 24.431 | 23.417 | [5 empty] | 0,20 | 19 | 4.0167 |  | 233743 | -0.0015 | 0.0028
2023-08-06 | 23:38 | MP | Ti3 |  | sMATF LED pri | 20.4588 | 24.1438 | 22.2773 | 23.7641 | 24.431 | 23.417 | [5 empty] | 1,76 |  |  |  | 233841 | 0.0007 | 0.0006 | 809.3 | 601.417
2023-08-06 |  | EB | Ti3 | Y | sMATF mirror | [13 empty] | 4.0167 |  |  | -0.00185 | 0.00265 | 0.021 | 0.314
2023-08-06 |  | EB | Ti3 | Y | sMATF LED pri | [13 empty] | 4e-06 |  |  |  |  | 809.2895 | 601.26 | Average of sMATF measurments
2023-08-07 |  | EB | INITIAL |  | sMATF mirror | 20.4588 | 24.1438 | 22.2773 | 23.7641 | 24.431 | 23.417 | [5 empty] | 0,16 |  |  |  |  | 0.0003 | 0.0025 |  |  | ADM int1 temp=295K, press=855 mbar(637.5 Torr torr)
2023-08-07 | 22:06 | EB | INITIAL |  | sMATF mirror | 20.4588 | 24.1438 | 22.2773 | 23.7641 | 24.431 | 23.417 | [5 empty] | 0,16 | 18 | 4.0167 |  | 40838 | 0.0017 | 0.0035
2023-08-07 | 21:33 | EB | INITIAL |  | sMATF LED pri | 20.4588 | 24.1438 | 22.2773 | 23.7641 | 24.431 | 23.417 | [5 empty] | 1,76 |  |  |  | 33254 |  |  | 809.059 | 602.648
2023-08-07 | 22:13 | EB | INITIAL |  | sMATF LED pri | 20.4588 | 24.1438 | 22.2773 | 23.7641 | 24.431 | 23.417 | [5 empty] | 1,76 |  |  |  | 41207 |  |  | 809.312 | 602.936
2023-08-07 | 22:28 | EB | INITIAL |  | sMATF retro | 20.4588 | 24.1438 | 22.2773 | 23.7641 | 24.15 | 23.75 | [5 empty] | 0,20 | 12 | 4.0272
2023-08-07 | 22:29 | EB | INITIAL |  | sMATF mirror | 20.4588 | 24.1438 | 22.2773 | 23.7641 | 24.431 | 23.417 | [5 empty] | 0,16 | 15 | 4.0167 |  | 43135 | 0.0031 | 0.0042
2023-08-07 | 22:33 | EB | INITIAL |  | sMATF LED pri | 20.4588 | 24.1438 | 22.2773 | 23.7641 | 24.431 | 23.417 | [5 empty] | 1,76 | [6 empty] | 809.65 | 603.368
2023-08-07 | 22:49 | EB | INITIAL |  | sMATF mirror | [11 empty] | 0,16 |  |  |  |  | 0.0031 | 0.0046
2023-08-07 | 22:51 | EB | INITIAL |  | sMATF LED pri | [11 empty] | 1,76 |  |  |  | 44916 |  |  | 809.818 | 603.287
2023-08-07 | 22:58 | EB | INITIAL |  | sMATF mirror | [11 empty] | 0,16 |  |  |  | 45207 | 0.0031 | 0.0049
2023-08-07 | 23:01 | EB | INITIAL |  | sMATF LED pri | [11 empty] | 1,76 |  |  |  | 45416 |  |  | 809.88 | 603.148
2023-08-07 | 23:06 | EB | INITIAL |  | sMATF mirror | [11 empty] | 0,16 |  |  |  | 50601 | 0.0031 | 0.0049
2023-08-07 | 23:14 | EB | INITIAL |  | sMATF LED pri | [11 empty] | 1,76 |  |  |  | 51220 |  |  | 809.65 | 603.33
2023-08-07 | 23:29 | EB | INITIAL |  | sMATF LED pri | [11 empty] | 1,76 | [6 empty] | 809.85 | 603.3
2023-08-07 | 0:10 | EB | INITIAL |  | sMATF LED pri | [11 empty] | 1,76 |  |  |  | 61258 |  |  | 809.907 | 603.406
2023-08-07 | 0:15 | EB | INITIAL |  | sMATF mirror | [11 empty] | 0,17 |  |  |  |  | 0.0031 | 0.0046
2023-08-07 | 0:40 | EB | nulling |  | sMATF mirror | 20.4468 | 24.1638 | 22.3059 | 23.7498 | 24.431 | 23.417 | [5 empty] | 0,17 |  |  |  | 064016 | -0.0004 | 0.0001
2023-08-07 |  | EB,MP | nulling |  | sMATF LED pri | 20.4468 | 24.1638 | 22.3059 | 23.7498 | 24.431 | 23.417 | [5 empty] | 1,76 | [6 empty] | 809.101 | 600.165
2023-08-07 | 0:57 | EB,MP | NULLED |  | sMATF LED pri | 20.4468 | 24.1638 | 22.3059 | 23.7498 | 24.431 | 23.417 | [5 empty] | 1,76 |  |  |  | 065745 |  |  | 808.481 | 602.166
2023-08-07 | 0:59 | EB,MP | NULLED |  | sMATF mirror | 20.4468 | 24.1638 | 22.3059 | 23.7498 | 24.431 | 23.417 | [5 empty] | 0,17 | 18 | 4.0167 |  | 065929 | 0.0003 | 0.0015
2023-08-07 | 1:27 | EB,MP | Settling |  | sMATF mirror | 20.4468 | 24.1638 | 22.3059 | 23.7498 | 24.431 | 23.417 | [5 empty] | 0,17 |  |  |  | 072735 | -0.0022 | 0.0028
2023-08-07 | 1:37 | EB,MP | Settling |  | sMATF LED pri | 20.4484 | 24.1651 | 22.3059 | 23.7498 | 24.431 | 23.417 | [5 empty] | 1,76 |  |  |  | 073728 |  |  | 808.13 | 602.608 | (Camera reported error; restarted Labview)
2023-08-07 | 1:40 | EB,MP | Settling |  | sMATF mirror | 20.4484 | 24.1651 | 22.3059 | 23.7498 | 24.431 | 23.417 | [5 empty] | 0,14 |  |  |  | 074041 | -0.0004 | 0.0022
2023-08-07 | 2:00 | EB,MP | Settling |  | sMATF mirror | 20.4484 | 24.1651 | 22.3059 | 23.7498 | 24.431 | 23.417 | [5 empty] | 0,17 |  |  |  | 080037 | -0.0022 | 0.0035
2023-08-07 | 2:01 | EB,MP | Settling |  | sMATF LED pri | 20.4484 | 24.1651 | 22.3059 | 23.7498 | 24.431 | 23.417 | [5 empty] | 1,76 |  |  |  | 080142 |  |  | 810.87 | 600.086
2023-08-07 | 2:16 | EB,MP | Settling |  | sMATF LED pri | 20.4484 | 24.1651 | 22.3059 | 23.7498 | 24.431 | 23.417 | [5 empty] | 1,76 |  |  |  | 081627 |  |  | 811.08 | 599.629
2023-08-07 | 2:18 | EB,MP | Settling |  | sMATF mirror | 20.4484 | 24.1651 | 22.3059 | 23.7498 | 24.431 | 23.417 | [5 empty] | 0,17 |  |  |  | 081817 | -0.0025 | 0.0039
2023-08-07 | 2:31 | EB,MP | Settling |  | sMATF mirror | 20.4484 | 24.1651 | 22.3059 | 23.7498 | 24.431 | 23.417 | [5 empty] | 0,17 |  |  |  | 083132 | -0.0028 | 0.0042
2023-08-07 | 2:33 | EB,MP | Settling |  | sMATF LED pri | 20.4484 | 24.1651 | 22.3059 | 23.7498 | 24.431 | 23.417 | [5 empty] | 0,17 |  |  |  | 083328 |  |  | 811.15 | 599.87
2023-08-07 | 2:46 | EB,MP | Settling |  | sMATF LED pri | 20.4484 | 24.1651 | 22.3059 | 23.7498 | 24.431 | 23.417 | [5 empty] | 1,76 |  |  |  | 084617 |  |  | 811.109 | 599.932
2023-08-07 | 2:48 | EB,MP | Settling |  | sMATF mirror | 20.4484 | 24.1651 | 22.3059 | 23.7498 | 24.431 | 23.417 | [5 empty] | 0,17 |  |  |  | 084820 | -0.0022 | 0.0042
2023-08-07 | 3:00 | EB,MP | Settling |  | sMATF LED pri | 20.4484 | 24.1651 | 22.3059 | 23.7498 | 24.431 | 23.417 | [5 empty] | 0,17 |  |  |  | 090057 |  |  | 811.159 | 599.914
2023-08-07 | 3:02 | EB,MP | Settling |  | sMATF mirror | 20.4484 | 24.1651 | 22.3059 | 23.7498 | 24.431 | 23.417 | [5 empty] | 1134 | 31 | 4.0167 |  | 090257 | -0.0022 | 0.0039 | -0.0022
2023-08-07 | 3:17 | EB,MP | TI3 v2 | Y | sMATF retro | 20.4484 | 24.1651 | 22.3059 | 23.7498 | 24.15 | 23.75 | [6 empty] | 15 | 4.0272 |  |  |  |  | 1e-06
2023-08-07 | 3:16 | EB,MP | TI3 v2 |  | sMATF mirror | 20.4484 | 24.1651 | 22.3059 | 23.7498 | 24.431 | 23.417 | [5 empty] | 0,17 | 15 | 4.0167 |  | 091657 | -0.0022 | 0.0042 |  |  | Pressure = 636..6 torr
2023-08-07 | 3:19 | EB,MP | TI3 v2 |  | sMATF LED pri | 20.4484 | 24.1651 | 22.3059 | 23.7498 | 24.431 | 23.417 | [5 empty] | 0,17 |  |  |  | 091903 |  |  | 811.311 | 599.932 | Temp = 295
2023-08-07 | 3:38 | EB,MP | Ti3 v2 | Y | PATB mirror | 20.4484 | 24.1651 | 22.3059 | 23.7498 | 24.431 | 23.405 | 101.524 | 125.748 | -0.002 | 0.0788 | 0.09 | 0,17 | 48 | 3.4693 |  | 093834 | -0.0035 | 0.0028 |  |  | No nulling required
2023-08-07 | 3:57 | EB,MP | Ti3 v2 | Y | PATB LED pri | 20.4484 | 24.1651 | 22.3059 | 23.7498 | 24.431 | 23.405 | 88.486 | 109.307 | -0.002 | 0.079 | 0.09 | 1130 |  | -1.1e-05 |  | 095712 |  |  | 811.254 | 600.059
2023-08-07 | 4:03 | EB,MP | Ti3 v2 | Y | PATA mirror | 20.4484 | 24.1651 | 22.3059 | 23.7498 | 24.431 | 23.405 | -76.125 | 125.703 | -0.002 | 0.0517 | 0.0624 | 0,20 | 34 | 3.4693 |  | 100323 | -0.0032 | 0.0015
2023-08-07 | 4:10 | EB,MP | Ti3 v2 | Y | PATA LED pri | 20.4484 | 24.1651 | 22.3059 | 23.7498 | 24.431 | 23.405 | -60.651 | 109.304 | 0 | 0.0518 | 0.0624 | 1130 |  | 1.1e-05 |  | 101044 |  |  | 810.954 | 599.853
2023-08-07 | 4:29 | EB,MP | Ti3 v2 |  | sMATF mirror | 20.4484 | 24.1651 | 22.3059 | 23.7498 | 24.431 | 23.417 | [5 empty] | 0,20 | 19 | 4.0167 |  | 102939 | -0.0028 | 0.0042 |  |  | Range with RM3 adjusted to PATB = 4.016746
2023-08-07 | 4:31 | EB,MP | Ti3 v2 |  | sMATF LED pri | 20.4484 | 24.1651 | 22.3059 | 23.7498 | 24.431 | 23.417 | [5 empty] | 1,76 |  |  |  | 103134 |  |  | 811.25 | 599.92
2023-08-07 |  | EB | Ti3 v2 | Y | sMATF mirror | [13 empty] | 4.0167 |  |  | -0.0025 | 0.00405
2023-08-07 |  | EB | Ti3 v2 | Y | sMATF LED pri | [18 empty] | 811.2805 | 599.926 | Average of sMATF measuremets
2023-08-08 | 14:36 | MP | Ti3 v3 |  | sMATF mirror | 20.4484 | 24.1651 | 22.3059 | 23.7498 | 24.431 | 23.417 | [5 empty] | 0,20 | 35 | 4.0167 |  | 203658 | -0.0015 | 0.0042 |  |  | NO NULLING REQUIRED
2023-08-08 | 14:43 | MP | Ti3 v3 |  | sMATF LED pri | 20.4484 | 24.1651 | 22.3059 | 23.7498 | 24.431 | 23.417 | [5 empty] | 1,76 |  | -1.55e-05 |  | 204305 | 0.0013 | 0 | 811.57 | 600.944 | Drift (1 mdeg, 1 pix) is within limit (<=5 mdeg, <= 1 pix)
2023-08-08 | 15:17 | MP | Ti3 v3 | Y | sMATF retro | 20.4484 | 24.1651 | 22.3059 | 23.7498 | 24.15 | 23.75 | [5 empty] | 0,20 | 34 | 4.0272 |  |  |  |  | 0.2895 | 1.018 | Temp=295, Pressure=634.5 torr
2023-08-08 | 15:23 | MP | Ti3 v3 |  | sMATF mirror | 20.4484 | 24.1651 | 22.3059 | 23.7498 | 24.431 | 23.417 | [5 empty] | 0,20 | 35 | 4.0167 |  | 212347 | -0.0015 | 0.0042
2023-08-08 | 15:22 | MP | Ti3 v3 |  | sMATF LED pri | 20.4484 | 24.1651 | 22.3059 | 23.7498 | 24.431 | 23.417 | [5 empty] | 1,76 |  |  |  | 212221 |  |  | 811.673 | 600.789
2023-08-08 | 15:23 | MP | Ti3 v3 | Y | PATB mirror | 20.4484 | 24.1651 | 22.3059 | 23.7498 | 24.431 | 23.417 | 101.524 | 125.748 | -0.002 | 0.0788 | 0.09 | 0,20 | 25 | 3.4693 |  | 212347 | -0.0032 | 0.0025
2023-08-08 | 15:44 | MP | Ti3 v3 | Y | PATB LED pri | 20.4484 | 24.1651 | 22.3059 | 23.7498 | 24.431 | 23.417 | 88.486 | 109.307 | -0.002 | 0.079 | 0.09 | 1130 |  | -1.1e-05 |  | 214417 | -0.0017 | -0.0017 | 811.378 | 601.07
2023-08-08 | 15:52 | MP | Ti3 v3 |  | PATA mirror | 20.4484 | 24.1651 | 22.3059 | 23.7498 | 24.431 | 23.417 | -76.125 | 125.703 | -0.002 | 0.0517 | 0.0624 | 0,20 |  |  |  | 215246 | -0.0028 | 0.0022 | -0.295 | 0.281
2023-08-08 | 15:59 | MP | Ti3 v3 | Y | PATA mirror | 20.4484 | 24.1651 | 22.3059 | 23.7498 | 24.431 | 23.417 | -76.125 | 125.703 | -0.002 | 0.0537 | 0.0634 | 0,20 | 46 | 3.4693 |  | 215948 | -0.0015 | 0.0042
2023-08-08 | 16:12 | MP | Ti3 v3 | Y | PATA LED pri | 20.4484 | 24.1651 | 22.3059 | 23.7498 | 24.431 | 23.417 | -60.651 | 109.304 | 0 | 0.0518 | 0.0624 | 1130 |  | -4.7e-05 |  | 221231 | 0 | 0 | 811.37 | 600.29
2023-08-08 | 16:40 | MP | Ti3 v3 |  | sMATF mirror | 20.4484 | 24.1651 | 22.3059 | 23.7498 | 24.431 | 23.417 | [5 empty] | 0,20 | 24 | 4.0167 |  | 224002 | -0.0022 | 0.0046 |  | -0.499
2023-08-08 | 16:44 | MP | Ti3 v3 |  | sMATF LED pri | 20.4484 | 24.1651 | 22.3059 | 23.7498 | 24.431 | 23.417 | [5 empty] | 1,76 |  | -8e-06 |  | 224412 | -0.0007 | 0.0004 | 811.746 | 600.74
2023-08-08 |  | MP | Ti3 v3 | Y | sMATF mirror | [13 empty] | 4.0167 |  |  | -0.00185 | 0.0044
2023-08-08 |  | MP | Ti3 v3 | Y | sMATF LED pri | [18 empty] | 811.7095 | 600.7645 | Average of sMATF measuremets
2023-08-10 | 8:52 | MP | Aperture Wheel Test | [6 empty] | 24.436 | 23.387 | [5 empty] | 0,25 |  |  |  | 145226 | -0.0435 | 0.0108 |  |  | AW @ Flat Mirror center (-856,000 counts)
2023-08-10 |  | MP | Aperture Wheel Test | [6 empty] | 24.436 | 23.387 | [5 empty] | 0,25 | 47-70 | [7 empty] | Steered RMs while measuring but was unsuccessful getting range meas.. Too far afield gave a BROKEN_BEAM error.
2023-08-10 | 8:55 | MP | Aperture Wheel Test | [6 empty] | 24.436 | 23.387 | [5 empty] | 1119 |  |  |  | 145514 |  |  |  |  | circular aperture in focus
2023-08-10 | 8:58 | MP | Aperture Wheel Test | [6 empty] | 24.436 | 23.387 | [5 empty] | 1153 |  |  |  | 145816 |  |  |  |  | vertical post in focus
2023-08-10 | 9:01 | MP | Aperture Wheel Test | [6 empty] | 24.436 | 23.387 | [5 empty] | 2106 |  |  |  | 150107 | 0.5111 | 0.0115 |  |  | catseye
2023-08-10 | 9:04 | MP | Aperture Wheel Test | [6 empty] | 24.436 | 23.387 | [5 empty] | 2107 |  |  |  | 150419 | 0.511 | -0.3319 |  |  | catseye
2023-08-10 | 9:18 | MP | Aperture Wheel Test | [6 empty] | 24.436 | 23.387 | [5 empty] | 2260 |  |  |  | 151825 |  |  |  |  | Warning: Measurement laser is visible off ADM window
2023-08-10 | 9:31 | MP | Aperture Wheel Test | [6 empty] | 24.436 | 23.387 | [5 empty] | 0,25 |  |  |  | 153151 |  |  |  |  | Using alignment laser to find edge. Laser walks rightward as wheel is rotated negative.. -1,008,000 counts is -Mirror
2023-08-10 | 9:43 | MP | Aperture Wheel Test | [6 empty] | 24.436 | 23.387 | [5 empty] | 0,25 |  |  |  | 154322 |  |  |  |  | Using alignment laser to find edge. Laser walks leftward has wheel is rotated positive... -690,000 counts is -Mirror
2023-08-10 | 9:46 | MP | Aperture Wheel Test | [6 empty] | 24.436 | 23.387 | [5 empty] | 0,25 |  |  |  | 154642 | -0.0431 | -0.006 |  |  | Do Not Use this position (-849,000 counts)
2023-08-10 |  | MP | Aperture Wheel Test | [22 empty] | Oberved negative edge of open aperture using alignment laser while SORC was in STOW position. Position was similar to previously measured position (7/15). SORC was moved out of STOW position, thus, blocking alignment laser return. Test was stopped.
2023-08-12 | 11:55 | EB | INITIAL |  | sMATF mirror |  |  |  |  | 24.443 | 23.421 | [5 empty] | 0,22 | 17 | 4.0234 |  | 175556 | 0.0499 | 0.0133 |  |  | ADM int1 = 299.81 K, 636.5 Torr. RMs were turned off and so we have no absolute encoder readings, made small relative adjustments to RM3 to optimize range measurement.
2023-08-12 | 12:15 | EB | INITIAL |  | sMATF LED pri |  |  |  |  | 24.443 | 23.421 | [5 empty] | 1,76 |  |  |  | 181545 |  |  | 842.34 | 664.292
2023-08-12 | 12:32 | EB | after homing |  | sMATF mirror | 20.4484 | 24.1651 | 22.3059 | 23.7498 | 24.443 | 23.421 | [5 empty] | 0,19 |  |  |  | 183232 | 0.0395 | 0.0431 |  |  | Couldn't acquire enough signal to range off mirror
2023-08-12 | 12:28 | EB | after homing |  | sMATF LED pri | 20.4484 | 24.1651 | 22.3059 | 23.7498 | 24.443 | 23.421 | [5 empty] | 1,76 |  |  |  | 182809 |  |  | 868.165 | 653.046
2023-08-12 | 12:45 | EB | nulling |  | sMATF LED pri | 20.4909 | 24.1256 | 22.3059 | 23.7498 | 24.443 | 23.421 | [5 empty] | 1,76 |  |  |  | 184543 |  |  | 807.709 | 602.212
2023-08-12 | 12:48 | EB | nulling |  | sMATF mirror | 20.4909 | 24.1256 | 22.3059 | 23.7498 | 24.443 | 23.421 | [5 empty] | 0,19 |  |  |  | 184853 | -0.0195 | -0.0263
2023-08-12 | 13:01 | EB | nulling |  | sMATF LED pri | 20.6014 | 23.9616 | 22.0705 | 23.8532 | 24.443 | 23.421 | [5 empty] | 1,76 |  |  |  | 190118 |  |  | 807.675 | 602.346
2023-08-12 | 13:03 | EB | nulling |  | sMATF mirror | 20.6014 | 23.9616 | 22.0705 | 23.8532 | 24.443 | 23.421 | [5 empty] | 0,21 |  |  |  | 190357 | -0.0094 | -0.0103
2023-08-12 | 13:11 | EB | nulling |  | sMATF LED pri | 20.7074 | 23.8361 | 22.0705 | 23.8532 | 24.443 | 23.421 | [5 empty] | 1,76 |  |  |  | 191146 |  |  | 808.353 | 602.435
2023-08-12 | 13:17 | EB | nulling |  | sMATF mirror | 20.7074 | 23.8361 | 21.8861 | 23.9528 | 24.443 | 23.421 | [5 empty] | 0,19 |  |  |  | 191748 | -0.0004 | 0.0022
2023-08-12 | 13:21 | EB | nulling |  | sMATF LED pri | 20.7079 | 23.8361 | 21.8861 | 23.9528 | 24.443 | 23.421 | [5 empty] | 1,76 |  |  |  | 192148 |  |  | 807.939 | 602.056
2023-08-12 | 13:24 | EB | nulling |  | sMATF mirror | 20.7079 | 23.8361 | 21.8861 | 23.9528 | 24.443 | 23.421 | [5 empty] | 0,19 |  |  |  | 192418 | -0.0008 | 0.0008
2023-08-12 | 13:54 | EB | NULLED |  | sMATF LED pri | 20.7069 | 23.8351 | 21.8861 | 23.9528 | 24.443 | 23.421 | [5 empty] | 1,76 |  |  |  | 195437 |  |  | 807.697 | 602.044
2023-08-12 | 13:56 | EB | NULLED |  | sMATF mirror | 20.7069 | 23.8351 | 21.8861 | 23.9528 | 24.418 | 23.416 | [5 empty] | 0,18 | 18 | 4.0235 |  | 195635 | -0.0001 | 0.0001
2023-08-12 | 14:05 | EB | Ti4 | Y | sMATF retro | 20.7069 | 23.8351 | 21.8861 | 23.9528 | 24.15 | 23.75 | [6 empty] | 17 | 4.034
2023-08-12 | 14:08 | EB | Ti4 | Y | CHIPS retro | 20.7069 | 23.8351 | 21.8861 | 23.9528 | 24.63 | 23.18 | [6 empty] | 11 | 4.7649 | [6 empty] | PRIOR CHIPS retro measurements are actually sMATF retro measurements. Note to self: go fix
2023-08-12 | 14:11 | EB | settling |  | sMATF mirror | 20.7069 | 23.8351 | 21.8861 | 23.9528 | 24.428 | 23.42 | [5 empty] | 0,18 | 19 | 4.0235 |  | 201133 | -0.0015 | 0.0015
2023-08-12 | 14:17 | EB | settling |  | sMATF LED pri | 20.7069 | 23.8351 | 21.8861 | 23.9528 | 24.428 | 23.42 | [5 empty] | 1,76 |  |  |  | 201712 |  |  | 808.732 | 600.999
2023-08-12 | 14:25 | EB | settling |  | sMATF mirror | 20.7069 | 23.8351 | 21.8861 | 23.9528 | 24.428 | 23.42 | [5 empty] | 0,18 |  |  |  | 202516 | -0.0022 | 0.0015
2023-08-12 | 14:27 | EB | settling |  | sMATF LED pri | 20.7069 | 23.8351 | 21.8861 | 23.9528 | 24.428 | 23.42 | [5 empty] | 1,76 |  |  |  | 202753 |  |  | 808.945 | 600.499
2023-08-12 | 14:44 | EB | settling |  | sMATF mirror | 20.7069 | 23.8351 | 21.8861 | 23.9528 | 24.428 | 23.42 | [5 empty] | 0,18 |  |  |  | 204437 | -0.0018 | 0.0018
2023-08-12 | 14:46 | EB | settling |  | sMATF LED pri | 20.7069 | 23.8351 | 21.8861 | 23.9528 | 24.428 | 23.42 | [5 empty] | 1,76 |  |  |  | 204657 |  |  | 809.29 | 600.36
2023-08-12 | 15:05 | EB | settling |  | sMATF mirror | 20.7069 | 23.8351 | 21.8861 | 23.9528 | 24.428 | 23.42 | [5 empty] | 0,18 |  |  |  | 210521 | -0.0022 | 0.0018
2023-08-12 | 15:07 | EB | settling |  | sMATF LED pri | 20.7069 | 23.8351 | 21.8861 | 23.9528 | 24.428 | 23.42 | [5 empty] | 1,76 |  |  |  | 210733 |  |  | 809.43 | 600.25
2023-08-12 | 15:25 | EB | settling |  | sMATF mirror | 20.7069 | 23.8351 | 21.8861 | 23.9528 | 24.428 | 23.42 | [5 empty] | 0,18 | 27 | 4.0235 |  | 212536 | -0.0018 | 0.0018
2023-08-12 | 15:31 | EB | settling |  | sMATF LED pri | 20.7069 | 23.8351 | 21.8861 | 23.9528 | 24.428 | 23.42 | [5 empty] | 1,76 |  |  |  | 213137 |  |  | 809.36 | 600.3
2023-08-12 | 16:25 | EB | Ti4 |  | sMATF mirror | 20.7069 | 23.8351 | 21.8861 | 23.9528 | 24.428 | 23.42 | [5 empty] | 0,18 | 43 | 4.0235 |  | 222542 | -0.0018 | 0.0022
2023-08-12 | 16:28 | EB | Ti4 |  | sMATF LED pri | 20.7069 | 23.8351 | 21.8861 | 23.9528 | 24.428 | 23.42 | [5 empty] | 1,76 |  |  |  | 222827 |  |  | 809.3 | 600.55
2023-08-12 | 16:48 | EB | prenull |  | PATB mirror | [6 empty] | 101.523 | 125.75 | -0.003 | 0.0784 | 0.0904 | 0,18 |  |  |  | 224842 | -0.0334 | -0.0156
2023-08-12 | 16:58 | EB | Ti4 | Y | PATB mirror | [6 empty] | 101.523 | 125.75 | -0.003 | 0.1144 | 0.0868 | 0,18 | 88 | 3.474 |  | 225817 | -0.0022 | 0.0032
2023-08-12 | 17:09 | EB | prenull |  | PATB LED pri | [11 empty] | 1130 |  |  |  | 230938 | -0.0004 | 0.001 | 796.73 | 586.16
2023-08-12 | 17:14 | EB | Ti4 | Y | PATB LED pri | [6 empty] | 88.57 | 108.385 | -0.003 | 0.1145 | 0.0868 | 1130 |  |  |  | 231403 |  |  | 809.33 | 600.6
2023-08-12 | 17:21 | EB | prenull |  | PATA mirror | [6 empty] | -76.132 | 125.703 | -0.0001 | 0.0533 | 0.063 | 0,18 |  |  |  | 232155 | -0.0216 | -0.0347 | 0.03 | 0.05
2023-08-12 | 17:33 | EB | Ti4 | Y | PATA mirror | [6 empty] | -76.132 | 125.703 | -0.0001 | 0.0911 | 0.0819 | 0,18 | 90 | 3.474 |  | 233303 | -0.0025 | 0.0028
2023-08-12 | 17:44 | EB | prenull |  | PATA LED pri | [6 empty] | -60.65 | 109.305 | -0.002 | 0.0911 | 0.0819 | 1130 |  |  |  | 234446 | -0.0007 | 0.0006 | 795.4 | 586.84
2023-08-12 | 17:48 | EB | Ti4 | Y | PATA LED pri | [6 empty] | -60.491 | 108.368 | -0.002 | 0.0912 | 0.082 | 1130 |  |  |  | 234856 |  |  | 809.43 | 600.56
2023-08-12 | -0.24236 | EB | Ti4 |  | sMATF mirror | [11 empty] | 0,18 | 45 | 4.0235 |  | 001111 | -0.0025 | 0.0022 | 0.13 | 0.01
2023-08-12 | -0.23889 | EB | Ti4 |  | sMATF LED pri | [11 empty] | 1,76 |  |  |  | 001617 |  |  | 809.34 | 599.8
2023-08-12 |  | EB | Ti4 | Y | sMATF mirror | [13 empty] | 4.0235 |  |  | -0.00215 | 0.0022
2023-08-12 |  | EB | Ti4 | Y | sMATF LED pri | [18 empty] | 809.32 | 600.175
TVAC1 Pre Ambient
2023-09-02 | 9:14 | CRD |  |  |  | RMs homed per ADM functional
2023-09-03 | 13:01 | CRD | Aperture Wheel Test |  | Aperture mirror | 20.45 | 24.17 | 22.31 | 23.75 | 24.436 | 23.387 | [5 empty] | 2,94 |  |  |  | 190644 | 0.5117 | 0.0112 |  |  | image also named "FlatMirror_2-94Focus.fits"
2023-09-03 | 13:11 | CRD | Aperture Wheel Test |  | Aperture mirror | 20.45 | 24.17 | 22.31 | 23.75 | 24.436 | 23.387 | [5 empty] | 0,26 |  |  |  | 190956 | -0.0539 | 0.0035 |  |  | image also named "FlatMirror_0-26Focus.fits"
Pre-metrology MATF AC check
2023-09-06 | 10:26 | EB | MATC AC check |  | MATF mirror | [11 empty] | 0,26 |  |  |  | 162626 | -0.0126 | -0.0222 |  |  | Starting temp of ADM Int1 is 291.7K
2023-09-06 | 10:36 | EB | MATC AC check |  | MATF LED pri | [11 empty] | 1,72 |  |  |  | 163622 |  |  | 833.4 | 663.467 | Central LED spot is VERY faint. Bottom left crosshair line is the only one visible, and it is very bright
2023-09-06 | 10:53 | EB | MATC AC check |  | MATF LED red | [11 empty] | 1,72 |  |  |  | 165321 |  |  | 834.8 | 664.426 | Upper right crosshair line is the only one visible
Pre-metrology MATF AC check
2023-09-07 | 9:43 | EB | After rough nulling |  | MATF mirror | 21.392 | 23.523 | 21.4035 | 24.5943 | 24.426 | 23.423 | [5 empty] | 0,21 | 19 | 4.0208 | 636.2 | 154058 | -0.0001 | 0.0042 |  |  | ADMint1 = 295.1K
2023-09-07 | 9:44 | EB | After rough nulling |  | MATF LED red | 21.392 | 23.523 | 21.4035 | 24.5943 | 24.426 | 23.423 | [5 empty] | 1,76 |  |  |  | 154448 |  |  | 807.69 | 601.64
2023-09-07 | 10:13 | EB | pre-null |  | PATB mirror | 21.392 | 23.523 | 21.4035 | 24.5943 | 24.426 | 23.423 | 101.522 | 125.755 | -0.001 | 0.114 | 0.0869 | 0,21 |  |  |  | 161331 | 0.0225 | 0.0549
2023-09-07 | 10:19 | EB | After rough nulling |  | PATB mirror | 21.392 | 23.523 | 21.4035 | 24.5943 | 24.426 | 23.423 | 101.522 | 125.755 | -0.001 | 0.0667 | 0.0576 | 0,19 | 31 | 3.4735 | 636.3 | 161932 | 0.0003 | 0.0035
2023-09-07 | 10:30 | EB | pre-null |  | PATB LED pri | 21.392 | 23.523 | 21.4035 | 24.5943 | 24.426 | 23.423 | 88.571 | 108.385 | 0 | 0.0667 | 0.0576 | 1134 |  |  |  | 163008 |  |  | 815.4 | 563.15
2023-09-07 | 10:34 | EB | After rough nulling |  | PATB LED pri | 21.392 | 23.523 | 21.4035 | 24.5943 | 24.426 | 23.423 | 87.221 | 107.075 | 0 | 0.0667 | 0.0576 | 1134 |  |  |  | 163414 |  |  | 807.24 | 601.57
2023-09-07 | 11:25 | EB | Beginning excursions |  | MATF LED red | 21.392 | 23.523 | 21.4035 | 24.5943 | 24.426 | 23.423 | [5 empty] | 1,73 |  |  |  | 172556 |  |  | 807.814 | 601.323
2023-09-07 | 11:47 | EB | Beginning excursions |  | MATF mirror | 21.792 | 23.523 | 21.4035 | 24.9943 | 24.426 | 23.423 | [5 empty] | 0,18 |  |  |  | 174703 | 0.094 | 0.0077
2023-09-07 | 11:54 | EB |  |  | MATF mirror | 21.792 | 23.523 | 21.4035 | 24.8943 | 24.426 | 23.423 | [5 empty] | 0,18 |  |  |  | 175440 | -0.0671 | 0.0046
2023-09-07 | 12:13 | EB |  |  | MATF LED red | 21.792 | 23.523 | 21.4035 | 25 | 24.426 | 23.423 | [5 empty] | 1,73 |  |  |  | 181308 |  |  | 809.26 | 655.61
2023-09-07 |  | EB |  |  | MATF mirror | 21.792 | 23.523 | 21.4035 | 25 | 24.426 | 23.423 | [5 empty] | 0,18 |  |  |  |  | 0.1006 | 0.0077
2023-09-07 | 12:23 | EB |  |  | MATF mirror | 21.862 | 23.523 | 21.4035 | 25 | 24.426 | 23.423 | [9 empty] | 182343 | -0.0008 | 0.0084
2023-09-07 | 12:26 | EB |  |  | MATF LED red | 21.862 | 23.523 | 21.4035 | 25 | 24.426 | 23.423 | [5 empty] | 1,73 |  |  |  | 182658 |  |  | 810.3 | 569.4
2023-09-07 | 12:32 | EB | Going back to baseline |  | MATF LED red | 21.392 | 23.523 | 21.4035 | 24.5943 | 24.426 | 23.423 | [5 empty] | 1,76 |  |  |  | 183241 |  |  | 812.389 | 592.597
2023-09-07 | 12:41 | EB | After re-homing |  | MATF LED red | 21.392 | 23.523 | 21.404 | 24.594 | 24.426 | 23.423 | [5 empty] | 1,76 |  |  |  | 184144 |  |  | 799.596 | 626.351
2023-09-07 | 12:58 | EB |  |  | MATF mirror | 21.392 | 23.523 | 21.404 | 24.594 | 24.426 | 23.423 | [5 empty] | 0,21 |  |  |  | 185808 | 0.0291 | -0.0037
2023-09-07 | 12:46 | EB |  |  | MATF LED red | 20.992 | 23.523 | 21.404 | 24.594 | 24.426 | 23.423 | [5 empty] | 1,76 |  |  |  | 184629 |  |  | 810.75 | 1142
2023-09-07 | 12:59 | EB |  |  | MATF mirror | 20.992 | 23.523 | 21.404 | 24.194 | 24.426 | 23.423 | [5 empty] | 0,21 |  |  |  | 185956 | -0.0619 | -0.0062
2023-09-07 | 13:03 | EB |  |  | MATF mirror | 20.952 | 23.523 | 21.404 | 24.194 | 24.426 | 23.423 | [5 empty] | 0,21 |  |  |  | 190335 | 0.0006 | -0.0058
2023-09-07 | 13:06 | EB |  |  | MATF LED red | 20.952 | 23.523 | 21.404 | 24.194 | 24.426 | 23.423 | [5 empty] | 1,76 |  |  |  | 190600 |  |  | 798.6 | 637.7
2023-09-07 | 13:12 | EB | Back to baseline |  | MATF LED red | 21.392 | 23.523 | 21.4035 | 24.5943 | 24.426 | 23.423 | [5 empty] | 1,76 |  |  |  | 191255 |  |  | 799.7 | 631.85
2023-09-07 | 13:26 | EB | After re-nulling |  | MATF LED red | 21.414 | 23.528 | 21.404 | 24.594 | 24.426 | 23.423 | [5 empty] | 1,76 |  |  |  | 192600 |  |  | 806.48 | 599.475
2023-09-07 | 13:16 | EB |  |  | MATF mirror | 21.414 | 23.528 | 21.404 | 24.594 | 24.426 | 23.423 | [5 empty] | 0,21 |  |  |  | 191622 | -0.0004 | 0.0028
2023-09-07 | 14:30 | EB | Monitoring |  | MATF LED red | 21.414 | 23.528 | 21.404 | 24.594 | 24.426 | 23.423 | [5 empty] | 1,76 |  |  |  | 203005 |  |  | 807.18 | 597.92
2023-09-07 | 14:34 | EB | Monitoring |  | MATF mirror | 21.414 | 23.528 | 21.404 | 24.594 | 24.426 | 23.423 | [5 empty] | 0,23 |  |  |  | 203404 | -0.0046 | 0.0042
2023-09-07 | 15:46 | EB |  |  | Other (specify in notes) | 21.414 | 23.528 | 21.404 | 24.594 | 24.426 | 23.423 | [5 empty] | 1286 |  |  |  | 214624 |  |  |  |  | transfer flat cat's eye
2023-09-07 | 15:52 | EB |  |  | Other (specify in notes) | 21.414 | 23.528 | 21.404 | 24.594 | 24.426 | 23.423 | [5 empty] | 0,18 |  |  |  | 215205 | 0.0965 | -0.0683 |  |  | transfer flat @ infinity
2023-09-07 | 15:55 | EB |  |  | Other (specify in notes) | 21.414 | 23.528 | 21.404 | 24.594 | 24.426 | 23.423 | [9 empty] | 215512 | 0.1138 | -0.1062 |  |  | transfer flat @ infinity
2023-09-07 | 16:01 | EB |  |  | Other (specify in notes) | 21.414 | 23.528 | 21.404 | 24.594 | 24.426 | 23.423 | [9 empty] | 220128 | 0.1524 | -0.1617 |  |  | transfer flat @ infinity
2023-09-07 | 16:06 | EB |  |  | Other (specify in notes) | 21.414 | 23.528 | 21.404 | 24.594 | 24.426 | 23.423 | [9 empty] | 220618 | 0.1006 | -0.0874 |  |  | transfer flat @ infinity
2023-09-07 | 16:08 | EB |  |  | Other (specify in notes) | 21.414 | 23.528 | 21.404 | 24.594 | 24.426 | 23.423 | [9 empty] | 220810 | 0.0756 | -0.0517 |  |  | transfer flat @ infinity
2023-09-07 | 16:09 | EB |  |  | Other (specify in notes) | 21.414 | 23.528 | 21.404 | 24.594 | 24.426 | 23.423 | [9 empty] | 220943 | 0.0159 | 0.0365 |  |  | transfer flat @ infinity
2023-09-07 | 16:20 | EB |  |  | Other (specify in notes) | 21.414 | 23.528 | 21.404 | 24.594 | 24.426 | 23.423 | [9 empty] | 222030 | 0.0145 | 0.0369 |  |  | transfer flat @ infinity
2023-09-07 | 16:27 | EB |  |  | Other (specify in notes) | 21.414 | 23.528 | 21.404 | 24.594 | 24.426 | 23.423 | [9 empty] | 222752 | 0.0142 | 0.0365 |  |  | transfer flat @ infinity
2023-09-07 | 17:18 | EB |  |  | Other (specify in notes) | 21.414 | 23.528 | 21.404 | 24.594 | 24.426 | 23.423 | [9 empty] | 231831 | 0.0138 | 0.0362 |  |  | transfer flat @ infinity
2023-09-07 | 17:40 | EB |  |  | Other (specify in notes) | 21.414 | 23.528 | 21.404 | 24.594 | 24.426 | 23.423 | [9 empty] | 234057 | 0.0135 | 0.0362 |  |  | transfer flat @ infinity
2023-09-07 | -0.24861 | EB |  |  | Other (specify in notes) | 21.414 | 23.528 | 21.404 | 24.594 | 24.426 | 23.423 | [9 empty] | 000244 | 0.0131 | 0.0358 |  |  | transfer flat @ infinity
2023-09-07 | -0.24167 | EB |  |  | Other (specify in notes) | 21.414 | 23.528 | 21.404 | 24.594 | 24.426 | 23.423 | [9 empty] | 001205 | 0.0131 | 0.0358 |  |  | transfer flat @ infinity
2023-09-07 | -0.23264 | EB |  |  | Other (specify in notes) | 21.414 | 23.528 | 21.404 | 24.594 | 24.426 | 23.423 | [9 empty] | 002513 | 0.0131 | 0.0355 |  |  | transfer flat @ infinity
2023-09-07 | -0.22569 | EB |  |  | Other (specify in notes) | 21.414 | 23.528 | 21.404 | 24.594 | 24.426 | 23.423 | [9 empty] | 003522 | 0.0131 | 0.0355 |  |  | transfer flat @ infinity
2023-09-07 | -0.22431 | EB |  |  | Other (specify in notes) | 21.414 | 23.528 | 21.404 | 24.594 | 24.426 | 23.423 | [9 empty] | 003715 | 0.0131 | 0.0355 |  |  | transfer flat @ infinity, with someone on TIF
2023-09-07 | -0.22222 | EB |  |  | MATF mirror | 21.414 | 23.528 | 21.404 | 24.594 | 24.426 | 23.423 | [9 empty] | 004029 | -0.0056 | 0.0032 |  |  | With someone on TIF
2023-09-07 | -0.22083 | EB |  |  | MATF mirror | 21.414 | 23.528 | 21.404 | 24.594 | 24.426 | 23.423 | [9 empty] | 004212 | -0.0053 | 0.0035 |  |  | w/o someone on TIF
Adjusting SALSA mirror
2023-09-08 | 8:50 | EB | after homing |  | MATF LED red | 21.414 | 23.528 | 21.404 | 24.594 | 24.426 | 23.423 | [5 empty] | 1,76 |  |  |  | 145008 |  |  | 802 | 569.63
2023-09-08 | 10:10 | EB | initial checkout |  | MATF mirror | 21.414 | 23.528 | 21.404 | 24.594 | 24.426 | 23.423 | [5 empty] | 0,18 |  |  |  | 161020 | 0.0034 | 0.0039
2023-09-08 | 10:12 | EB | initial checkout |  | MATF LED red | 21.414 | 23.528 | 21.404 | 24.594 | 24.426 | 23.423 | [5 empty] | 1,76 |  |  |  | 161207 |  |  | 806.7 | 604.9
2023-09-08 | 10:18 | EB | initial checkout |  | MATF LED pri | 21.414 | 23.528 | 21.404 | 24.594 | 24.426 | 23.423 | [5 empty] | 1,76 |  |  |  | 161825 |  |  | 805.35 | 605.95
2023-09-08 |  | EB | after 1st adjustment |  | MATF LED red | 20.788 | 23.57 | 21.443 | 23.858 | 24.426 | 23.423 | [5 empty] | 1,76 | [6 empty] | 807.25 | 601.81
2023-09-08 | 10:53 | EB | after 1st adjustment |  | MATF mirror | 20.788 | 23.57 | 21.443 | 23.858 | 24.426 | 23.423 | [5 empty] | 0,16 |  |  |  | 165326 | 0.001 | -0.0024
2023-09-08 | 11:11 | EB | after 2nd adjustment |  | MATF LED red | 20.469 | 23.99 | 22.197 | 23.457 | 24.426 | 23.423 | [5 empty] | 1,76 |  |  |  | 171143 |  |  | 807.52 | 603.21
2023-09-08 | 11:15 | EB | after 2nd adjustment |  | MATF mirror | 20.469 | 23.99 | 22.197 | 23.457 | 24.426 | 23.423 | [5 empty] | 0,16 |  |  |  | 171544 | 0.0034 | -0.009
2023-09-08 | 11:33 | EB | after 3rd adjustment |  | MATF LED red | 20.395 | 23.23 | 20.775 | 23.411 | 24.426 | 23.423 | [5 empty] | 1,76 |  |  |  | 173336 |  |  | 808.98 | 602.21
2023-09-08 | 11:36 | EB | after 3rd adjustment |  | MATF mirror | 20.395 | 23.23 | 20.775 | 23.411 | 24.435 | 23.433 | [5 empty] | 0,16 | 18 | 4.0206 | 638.5 | 173619 | -0.0028 | 0.0022
2023-09-08 | 13:23 | EB | Going back to 1st adjustment |  | MATF mirror | 20.788 | 23.57 | 21.443 | 23.858 | 24.42 | 23.433 | [5 empty] | 0,17 | 17 | 4.0208 | 637.7 | 192321 | 0.0013 | 0.0001
2023-09-08 | 13:26 | EB | Going back to 1st adjustment |  | MATF LED red | 20.788 | 23.57 | 21.443 | 23.858 | 24.42 | 23.433 | [5 empty] | 1,76 |  |  |  | 192615 |  |  | 808.17 | 603.077
Checking beam walk range
2023-09-08 | 13:40 | EB | When RM2B = 25 |  | MATF mirror | 22.1058 | 23.562 | 21.443 | 25 | 24.42 | 23.43 | [5 empty] | 0,17 |  |  |  | 194050 | -0.0046 | 0.0091
2023-09-08 | 13:43 | EB | When RM2B = 25 |  | MATF LED red | 22.1058 | 23.562 | 21.443 | 25 | 24.42 | 23.43 | [5 empty] | 1,76 |  |  |  | 194300 |  |  | 814.95 | 501.85 | Calculated beam walk range: | 6.009 | mm
2023-09-08 | 13:57 | EB | When RM1A=20 |  | MATF mirror | 20.048 | 23.58 | 21.443 | 23.218 | 24.42 | 23.43 | [5 empty] | 0,17 |  |  |  | 195737 | 0.0076 | 0.0084
|  |  | When RM1A=20 |  | MATF LED red | 20.048 | 23.58 | 21.443 | 23.218 | 24.42 | 23.43 | [12 empty] | 816 | 659.5 | Calculated beam walk range: | -3.45 | mm
Initial ADM data collect (rough nulling only)
2023-09-08 | 17:50 | EB | Rough ambient visual checks |  | MATF mirror | 20.793 | 23.555 | 21.443 | 23.858 | 24.425 | 23.423 | [5 empty] | 0,15 | 13 | 4.0207 | 636.3 | 235014 | 0.002 | 0.0018 |  |  | ADM Int1 = 298.3 K
2023-09-08 | 41:55 | EB | Rough ambient visual checks |  | MATF LED red | 20.793 | 23.555 | 21.443 | 23.858 | 24.425 | 23.423 | [5 empty] | 1,76 |  |  |  | 235512 |  |  | 808.06 | 603.34
2023-09-08 | 19:02 | EB | Rough ambient visual checks |  | PATB mirror | 20.793 | 23.555 | 21.443 | 23.858 | 24.43 | 23.423 | 101.52 | 125.759 | 0.001 | 0.0663 | 0.0575 | 0,18 | 15 | 3.4735 | 636.7 | 010233 | 0.0003 | 0.0008
2023-09-08 | 19:12 | EB | Rough ambient visual checks |  | PATB LED pri | 20.793 | 23.555 | 21.443 | 23.858 | 24.43 | 23.423 | 87.222 | 107.075 | 0.001 | 0.0664 | 0.0575 | 1129 |  |  |  | 011218 |  |  | 810.8 | 601.36
2023-09-08 | 19:19 | EB | Rough ambient visual checks |  | PATB LED red | 20.793 | 23.555 | 21.443 | 23.858 | 24.43 | 23.423 | 110.11 | 147.509 | 0.001 | 0.0661 | 0.0575 | 1129 |  |  |  | 011940 |  |  | 807.7 | 601.4
2023-09-08 | 19:26 | EB | Rough ambient visual checks |  | PATB retro pri | 20.793 | 23.555 | 21.443 | 23.858 | 24.43 | 23.423 | ?? | ?? | 0.001 | 0.0664 | 0.0575 |  | 13 | 3.4598 | 636.9
2023-09-08 | 19:35 | EB | Rough ambient visual checks |  | PATA mirror | 20.793 | 23.555 | 21.443 | 23.858 | 24.43 | 23.423 | -76.134 | 125.707 | -0.001 | 0.0459 | 0.0561 | 0,17 | 16 | 3.4736 | 637 | 013509 | 0.001 | 0.0039
2023-09-08 | 19:42 | EB | Rough ambient visual checks |  | PATA LED pri | 20.793 | 23.555 | 21.443 | 23.858 | 24.43 | 23.423 | -61.877 | 107.118 | -0.002 | 0.0459 | 0.0561 | 1129 |  |  |  | 014202 |  |  | 808.3 | 601.7
2023-09-08 | 19:49 | EB | Rough ambient visual checks |  | PATA LED red | 20.793 | 23.555 | 21.443 | 23.858 | 24.43 | 23.423 | -84.983 | 147.322 | 0 | 0.0459 | 0.0561 | 1129 |  |  |  | 014933 |  |  | 808.14 | 602.53
2023-09-08 | 19:53 | EB | Rough ambient visual checks |  | PATA retro pri | 20.793 | 23.555 | 21.443 | 23.858 | 24.43 | 23.423 | -72.079 | 98.977 | 0.001 | 0.0459 | 0.0561 |  | 22 | 3.4599 | 637.2 | 015358
|  |  | Ambient Vacuum | [22 empty] | ADM internal 1 = 296.4 K
2023-09-14 | 12:24 | CD | INITIAL |  | MATF mirror | 0 | 0 | 0 | 0 | 0 | 0 | ADM/H2 |  |  |  |  | 0,20 |  |  | 636.9 | 182430 | 0.0097 | 0.0021 |  |  | F9,F10,F8,F9,F10,F11
2023-09-14 | 12:28 | CD | INITIAL |  | MATF LED red | 0 | 0 | 0 | 0 | 0 | 0 | ADM/H2 |  |  |  |  | 1,73 |  |  |  | 182851 |  |  | 811.8 | 611.8 | F8,F7
2023-09-14 | 12:42 | CD | after homing |  | MATF mirror | 20.793 | 23.555 | 21.443 | 23.858 | 24.43 | 23.42 | ADM/H2 |  |  |  |  | 0,20 |  |  | 636.8 | 184241 | 0.0166 | 0.0133
2023-09-14 | 12:40 | CD | after homing |  | MATF LED red | 20.793 | 23.555 | 21.443 | 23.858 | 24.43 | 23.42 | ADM/H2 |  |  |  |  | 1,73 |  |  |  | 184036 |  |  | 821.2 | 617.7
2023-09-14 | 12:51 | CD | nulling |  | MATF mirror | 20.803 | 23.552 | 21.443 | 23.858 | 24.43 | 23.42 | ADM/H2 |  |  |  |  | 0,20 |  |  | 636.8 | 185103 | -0.0001 | -0.002
2023-09-14 | 12:49 | CD | nulling |  | MATF LED red | 20.803 | 23.552 | 21.443 | 23.858 | 24.43 | 23.42 | ADM/H2 |  |  |  |  | 1,73 |  |  |  | 184958 |  |  | 808 | 602 | I got the spot centered on the cross hair but didn't make an actual measurement (just said 808,602), wish I had because I think it was already nulled here!
2023-09-14 | 13:04 | CD | nulled |  | MATF mirror | 20.803 | 23.552 | 21.44 | 23.86 | 24.43 | 23.42 | ADM/H2 |  |  |  |  | 0,20 |  |  | 636.8 | 190409 | -0.0001 | -0.0024
2023-09-14 | 13:01 | CD | nulled |  | MATF LED red | 20.803 | 23.552 | 21.44 | 23.86 | 24.43 | 23.42 | ADM/H2 |  |  |  |  | 1,73 |  |  |  | 190158 |  |  | 808.2 | 603
2023-09-14 | 13:14 | CD | settling (measure every 15 mins) |  | MATF mirror | 20.803 | 23.552 | 21.44 | 23.86 | 24.43 | 23.42 | ADM/H2 |  |  |  |  | 0,20 |  |  | 636.8 | 191446 | -0.0015 | -0.0013 |  |  | 0.007
2023-09-14 | 13:16 | CD | settling |  | MATF LED red | 20.803 | 23.552 | 21.44 | 23.86 | 24.43 | 23.42 | ADM/H2 |  |  |  |  | 1,73 |  |  |  | 191611 |  |  | 809.1 | 601.5 | 6.8
2023-09-14 | 13:29 | CD | settling |  | MATF mirror | 20.803 | 23.552 | 21.44 | 23.86 | 24.43 | 23.42 | ADM/H2 |  |  |  |  | 0,20 |  |  | 636.6 | 192955 | -0.0015 | -0.0013 |  |  | 0
2023-09-14 | 13:31 | CD | settling |  | MATF LED red | 20.803 | 23.552 | 21.44 | 23.86 | 24.43 | 23.42 | ADM/H2 |  |  |  |  | 1,73 |  |  |  | 193137 |  |  | 809.3 | 601.1 | 1.6
2023-09-14 | 13:44 | CD | settling |  | MATF mirror | 20.803 | 23.552 | 21.44 | 23.86 | 24.43 | 23.42 | ADM/H2 |  |  |  |  | 0,20 |  |  | 636.5 | 194426 | -0.0018 | -0.0013 |  |  | 0.001
2023-09-14 | 13:45 | CD | settling |  | MATF LED red | 20.803 | 23.552 | 21.44 | 23.86 | 24.43 | 23.42 | ADM/H2 |  |  |  |  | 1,73 |  |  |  | 194546 |  |  | 809.4 | 601 | 0.4
2023-09-14 | 13:59 | CD | settling |  | MATF mirror | 20.803 | 23.552 | 21.44 | 23.86 | 24.43 | 23.42 | ADM/H2 |  |  |  |  | 0,20 |  |  | 636.5 | 195924 | -0.0022 | -0.0013 |  |  | 0.002
2023-09-14 | 14:00 | CD | settling |  | MATF LED red | 20.803 | 23.552 | 21.44 | 23.86 | 24.43 | 23.42 | ADM/H2 |  |  |  |  | 1,73 |  |  |  | 200040 |  |  | 809.4 | 600.7 | 1.2
2023-09-14 | 14:14 | CD | settling |  | MATF mirror | 20.803 | 23.552 | 21.44 | 23.86 | 24.43 | 23.42 | ADM/H2 |  |  |  |  | 0,20 |  |  | 636.4 | 201425 | -0.0022 | -0.0013 |  |  | 0
2023-09-14 | 14:16 | CD | settling |  | MATF LED red | 20.803 | 23.552 | 21.44 | 23.86 | 24.43 | 23.42 | ADM/H2 |  |  |  |  | 1,73 |  |  |  | 201600 |  |  | 809.4 | 600.5 | 0.8
2023-09-14 | 14:29 | CD | WFI1 |  | MATF mirror | 20.803 | 23.552 | 21.44 | 23.86 | 24.43 | 23.42 | ADM/H2 |  |  |  |  | 0,20 |  |  | 636.3 | 202925 | -0.0022 | -0.0017
2023-09-14 | 14:30 | CD | WFI1 |  | MATF LED red | 20.803 | 23.552 | 21.44 | 23.86 | 24.43 | 23.42 | ADM/H2 |  |  |  |  | 1,73 |  |  |  | 203044 |  |  | 809.3 | 600.4 | 10mA
2023-09-14 | 14:36 | CD | WFI1 |  | MATF LED pri | 20.803 | 23.552 | 21.44 | 23.86 | 24.43 | 23.42 | ADM/H2 |  |  |  |  | 1,73 |  |  |  | 203610 |  |  | 808.6 | 601.3 | 20mA
2023-09-14 | 15:06 | CD | WFI1 - trending |  | MATF mirror | 20.803 | 23.552 | 21.44 | 23.86 | 24.43 | 23.425 | ADM/H2 |  |  |  |  | 0,20 | 20 | 4.021 | 636.6 | 210954 | -0.0022 | -0.002 |  |  | got images with guide and IR beam: 210348 210508
2023-09-14 | 15:12 | CD | WFI1 - trending |  | MATF LED red | 20.803 | 23.552 | 21.44 | 23.86 | 24.43 | 23.425 | ADM/H2 |  |  |  |  | 1,73 |  |  |  | 211257 |  |  | 808.7 | 600.7
2023-09-14 | 16:15 | CD | WFI1 - trending |  | MATF mirror | 20.803 | 23.552 | 21.44 | 23.86 | 24.43 | 23.425 | ADM/H2 |  |  |  |  | 0,20 | 13 | 4.0211 | 637.3 | 221521 | -0.0035 | -0.0013 |  |  | 637.3 torr
2023-09-14 | 16:26 | CD | WFI1 - trending |  | MATF LED red | 20.803 | 23.552 | 21.44 | 23.86 | 24.43 | 23.425 | ADM/H2 |  |  |  |  | 1,73 |  |  |  | 222608 |  |  | 809.39 | 599.55
2023-09-14 | 16:46 | PM | WFI1 |  | MATF mirror | 20.803 | 23.552 | 21.44 | 23.86 | 24.43 | 23.425 | ADM/H2 |  |  |  |  | 0,20 | 14 | 4.0211 | 637.4 | 224636 | -0.0035 | -0.002
2023-09-14 | 16:44 | PM | WFI1 |  | MATF LED red | 20.803 | 23.552 | 21.44 | 23.86 | 24.43 | 23.425 | ADM/H2 |  |  |  |  | 1,73 |  |  |  | 224409 |  |  | 809.415 | 599.43
2023-09-14 | 17:06 | PM | WFI1 |  | MATF mirror | 20.803 | 23.552 | 21.44 | 23.86 | 24.43 | 23.425 | ADM/H2 |  |  |  |  | 0,20 | 31 | 4.0211 | 637.6 | 230638 | -0.0032 | -0.0024
2023-09-14 | 17:10 | PM | WFI1 |  | MATF LED red | 20.803 | 23.552 | 21.44 | 23.86 | 24.43 | 23.425 | ADM/H2 |  |  |  |  | 1,73 |  | 4.0211 |  | 231054 |  |  | 808.77 | 599.81
2023-09-14 | 17:21 | PM | WFI1 |  | Aperture mirror | 20.803 | 23.552 | 21.44 | 23.86 | 24.43 | 23.425 | ADM/H2 |  |  |  |  | 0,20 |  |  |  | 232159 | Error 380 | Error 380 |  |  | LES analysis: runtime Error 380; unable to analyze
2023-09-14 | 17:40 | PM | WFI1 |  | Aperture mirror | 20.803 | 23.552 | 21.44 | 23.86 | 24.43 | 23.425 | ADM/H2 |  |  |  |  | 0,22 |  |  |  | 234028 |  |  |  |  | LES analysis: runtime Error 380; unable to analyze
2023-09-14 | 18:02 | PM | WFI1 |  | PATB mirror | 20.803 | 23.552 | 21.44 | 23.86 | 24.43 | 23.425 | 101.521 | 125.762 | -0.001 | 0.0659 | 0.0574 | 0,20 |  |  |  | 000243 | -0.0035 | -0.0076
2023-09-14 | 18:10 | PM | WFI1 | Y | PATB mirror | 20.803 | 23.552 | 21.44 | 23.86 | 24.43 | 23.425 | 101.521 | 125.762 | -0.001 | 0.0688 | 0.0612 | 0,20 | 90 | 3.4738 | 638.2 | 001034 | -0.0032 | -0.0024
2023-09-14 | 18:26 | PM | WFI1 |  | PATB LED pri | 20.803 | 23.552 | 21.44 | 23.86 | 24.43 | 23.425 |  |  | -0.001 | 0.069 | 0.0613 | 1132 |  |  |  | 002637 |  |  | 809.72 | 599.32
2023-09-14 | 18:27 | PM | WFI1 | Y | PATB LED pri | 20.803 | 23.552 | 21.44 | 23.86 | 24.43 | 23.425 | 87.188 | 107.072 | -0.001 | 0.069 | 0.0613 | 1132 |  |  |  | 002757 |  |  | 809.2 | 599.9
2023-09-14 | 18:39 | PM | WFI1 |  | PATB LED red | 20.803 | 23.552 | 21.44 | 23.86 | 24.43 | 23.425 | 110.109 | 147.508 | -0.001 | 0.0686 | 0.0613 | 1132 |  |  |  | 003933 |  |  | 806.82 | 599.4
2023-09-14 | 18:42 | PM | WFI1 | Y | PATB LED red | 20.803 | 23.552 | 21.44 | 23.86 | 24.43 | 23.425 | 110.185 | 147.431 | -0.001 | 0.0686 | 0.0613 | 1132 |  |  |  | 004221 |  |  | 809.11 | 599.81
2023-09-14 | 18:55 | PM | WFI1 |  | PATA mirror | 20.803 | 23.552 | 21.44 | 23.86 | 24.43 | 23.425 | -76.136 | 125.708 | -0.002 | 0.0458 | 0.0557 | 0,20 | 30 | 3.4738 | 638.5 | 005501 | -0.0035 | -0.001
2023-09-14 | 19:04 | PM | WFI1 | Y | PATA mirror | 20.803 | 23.552 | 21.44 | 23.86 | 24.43 | 23.425 | -76.136 | 125.708 | -0.002 | 0.0452 | 0.0544 | 0,20 | 34 | 3.4739 | 638.5 | 010403 | -0.0035 | -0.0027
2023-09-14 | 19:10 | PM | WFI1 |  | PATA LED pri | 20.803 | 23.552 | 21.44 | 23.86 | 24.43 | 23.425 | -61.879 | 107.118 | -0.003 | 0.0453 | 0.0544 | 1132 |  |  |  | 011046 |  |  | 808.92 | 593.68
2023-09-14 | 19:22 | PM | WFI1 | Y | PATA LED pri | 20.803 | 23.552 | 21.44 | 23.86 | 24.43 | 23.425 | -61.902 | 107.067 | -0.003 | 0.0453 | 0.0544 | 1132 |  |  |  | 012257 |  |  | 809 | 599.91
2023-09-14 | 19:30 | PM | WFI1 | Y | PATA LED red | 20.803 | 23.552 | 21.44 | 23.86 | 24.43 | 23.425 | -84.985 | 147.32 | 0 | 0.045 | 0.0544 | 1132 |  |  |  | 013008 |  |  | 808.95 | 599.79
2023-09-14 | 19:53 | PM | WFI1 |  | MATF mirror | 20.803 | 23.552 | 21.44 | 23.86 | 24.43 | 23.425 | ADM/H2 |  |  |  |  | 0,20 | 13 | 4.0211 | 638.5 | 015320 | -0.0049 | -0.0027
2023-09-14 | 19:59 | PM | WFI1 |  | MATF LED red | 20.803 | 23.552 | 21.44 | 23.86 | 24.43 | 23.425 | ADM/H2 |  |  |  |  | 1,73 |  | 7.1e-05 |  | 015908 |  |  | 809.3 | 598.73
2023-09-14 | 20:03 | PM | WFI1 |  | MATF LED pri | 20.803 | 23.552 | 21.44 | 23.86 | 24.43 | 23.425 | ADM/H2 |  |  |  |  | 1,73 |  |  |  | 020338 |  |  | 808.11 | 599.56
2023-09-14 |  | EB | WFI1 | Y | MATF mirror | [13 empty] | 4.0211 | 638.05 |  | -0.00405 | -0.00255 |  |  | averaging initial/closeout
2023-09-14 |  | EB | WFI1 | Y | MATF LED red | [18 empty] | 809.035 | 599.27 | averaging initial/closeout
Note on A1: [Threaded comment] Your version of Excel allows you to read this threaded comment; however, any edits to it will get removed if the file is opened in a newer version of Excel. Learn more: https://go.microsoft.com/fwlink/?linkid=870924 Comment: Try Ctrl+;
Note on B1: [Threaded comment] Your version of Excel allows you to read this threaded comment; however, any edits to it will get removed if the file is opened in a newer version of Excel. Learn more: https://go.microsoft.com/fwlink/?linkid=870924 Comment: Try Ctrl+Shift+;
Note on D1: [Threaded comment] Your version of Excel allows you to read this threaded comment; however, any edits to it will get removed if the file is opened in a newer version of Excel. Learn more: https://go.microsoft.com/fwlink/?linkid=870924 Comment: WP1 = Warm Plateau 1 CP2 = Cold Plateau 2, etc
Note on E1: [Threaded comment] Your version of Excel allows you to read this threaded comment; however, any edits to it will get removed if the file is opened in a newer version of Excel. Learn more: https://go.microsoft.com/fwlink/?linkid=870924 Comment: If you have to measure a target multiple times, mark the last/best one with a "Y"
Note on F48: [Threaded comment] Your version of Excel allows you to read this threaded comment; however, any edits to it will get removed if the file is opened in a newer version of Excel. Learn more: https://go.microsoft.com/fwlink/?linkid=870924 Comment: Aperture wheel electronics not yet functional
Note on AA451: [Threaded comment] Your version of Excel allows you to read this threaded comment; however, any edits to it will get removed if the file is opened in a newer version of Excel. Learn more: https://go.microsoft.com/fwlink/?linkid=870924 Comment: Range measurement temperature was not updated from previous hot measurement.
Note on AA452: [Threaded comment] Your version of Excel allows you to read this threaded comment; however, any edits to it will get removed if the file is opened in a newer version of Excel. Learn more: https://go.microsoft.com/fwlink/?linkid=870924 Comment: 11:58 After unsuccessful New 5DOF reboot, success restarting FARM by power cycling Agilent "A1A7" (in old 5DOF rack) in order (ch 1 OFF, ch 4 OFF, USB at font of new 5 DOF disconnected, wait for agilent power volt on ch to reach 0 then ch 4 ON, wait for agilent power volt on ch 1 < 60 mv (about 1 min) then ch1 ON, replug USB at front of new 5DOF. The success was first observed because the ch1 current changed to 1.0A instead of ~0.9A.
Note on AA459: [Threaded comment] Your version of Excel allows you to read this threaded comment; however, any edits to it will get removed if the file is opened in a newer version of Excel. Learn more: https://go.microsoft.com/fwlink/?linkid=870924 Comment: This is the second consecutive time (previous on 8/2) where small adjustment to RM1 seemed to restore our NULL alignment from a few rows above.
Note on T467: [Threaded comment] Your version of Excel allows you to read this threaded comment; however, any edits to it will get removed if the file is opened in a newer version of Excel. Learn more: https://go.microsoft.com/fwlink/?linkid=870924 Comment: (Surprised we had no range return and had to search for it. stdev on this range was much better compared to Nulled)
Note on B563: [Threaded comment] Your version of Excel allows you to read this threaded comment; however, any edits to it will get removed if the file is opened in a newer version of Excel. Learn more: https://go.microsoft.com/fwlink/?linkid=870924 Comment: I think these times are off by 2h?

## PMscr
THIS PAGE COPIED FROM THE OTHER TUNERS. | [8 empty] | INSTRUCTIONS: (First, update references in Nulling calculators)
sMATF RM Nulling | [9 empty] | RM2 overcorrection for AZ/EL (mm/deg) |  | Angle between Az and Rx (deg) | (rad.) | Scale (pix/mm)
Date | RM1( with rough-in formula, so the LED will stay near the center after big RM2 overcorrections) |  | RM2 (with pulling-down formula) |  | AC AZ (deg) | AC EL (deg) | X centroid (pix) | Y centroid (pix) | 1. Center sMATF pinhole (disable mirrors immed. after move) 2. Measure AutoCollimation Az & El | 2B/AZ | 2A/EL | 34.5 | 0.60214 | 20.9
| Target values --------> |  |  |  | 0 | 0 | 808 | 602 | 3. Enter Az/El, RM2 in Positon Log then move RM2 (disable mirrors immed. after move) | 10.6 | -17.9 | Do not change these!
>> | 20.803 | 23.552 | 21.44 | 23.86 | -0.0001 | -0.0024 | 808.2 | 603 | 4. You can try moving RM1 to Mark's RM1 | RM1a/pix | RM1b/pix
Pete's RM1: | 20.8049 | 23.5524 | 21.397 | 23.8611 |  |  |  |  | 5. Repeat steps 1-4 until nulled (Az/El < 0.001, Pix<1) | -0.0018524 | 0.00066474
Mark's RM1: | 20.7841 | 23.5036
PATB RxRy Tuner | [8 empty] | Update Target Az/El
Date | Rx (deg) | Ry (deg) | AC Az (deg) | AC EL (deg) | Target Az(deg) | Target El(deg) | Az Err | El Err | 1. Measure AC Az/El 2. Use script to move to new Rx,Ry
2023-08-03 | 0.0659 | 0.0574 | -0.0032 | -0.0024 | -0.0032 | -0.0024 | 0 | 0 | 3. Repeat steps 1&2 until Az,El < 0.005
Go here -> | 0.0659 | 0.0574
Update references in orange cells above, and set Target to post settle sMATF value
PATB dXdY Tuner | [8 empty] | Update Target Px/Py
Date | dX (mm) | dY (mm) | Px (pix) | Py(pix) | Target Px (pix) | Target Py (pix) | Px Err | Py Err | 1. Measure LED location 2. Use script to move X,Y
>> | 0 | 0 | 809.11 | 599.81 | 809.1 | 599.81 | 0.01 | 0 | 3. After move, check to make sure stages moved in right direction (correct sign if not)
Deltas -> | -0.00039432 | 0.00027101 | [6 empty] | 4. Repeat steps 1&2 until Px Err and Py Err < 1
Update references in orange cells above, and set Target to post settle sMATF value
7/25 HAD TO CHANGE DELTAS SIGN
8/6 HAD TO CHANGE DELTAS again
PATA RxRy Tuner |  |  |  |  | Use when at PATB & PATA positions
Date | Rx (deg) | Ry (deg) | AC Az (deg) | AC EL (deg) | Target Az(deg) | Target El(deg) | Az Err | El Err
>> | 0.0452 | 0.0544 | -0.0035 | -0.0027 | -0.0032 | -0.0024 | -0.0003 | -0.0003
Go here -> | 0.045617 | 0.054477
Update references in orange cells above, and set Target to post settle sMATF value
PATA dXdY Tuner
Date | dX (mm) | dY (mm) | Px (pix) | Py(pix) | Target Px (pix) | Target Py (pix) | Px Err | Py Err
>> | 0 | 0 | 808.95 | 599.79 | 809.1 | 599.81 | -0.15 | -0.02
Deltas -> | 0.0053728 | -0.0048538
Update references in orange cells above, and set Target to post settle sMATF value
8/2 HAD TO RESTORE DELTAS SIGN
| Nulling.. | Az (deg) | EL (deg) | Px (pix) | Py (pix)
| sMATF | -0.0001 | -0.0024 | 0.2 | 1
| PATA | 0 | 0 | 0.01 | 0
| PATB | -0.0003 | -0.0003 | -0.15 | -0.02

## Rm Null
Date | RM1( with rough-in formula, so the pinhole will stay near the center after big RM2 overcorrections) |  | RM2 (with pulling-down formula) |  | AC AZ (deg) | AC EL (deg) | X centroid (pix) | Y centroid (pix) |  |  | RM2 overcorrection for AZ/EL (mm/deg) |  | 0. Confirm that you have used RM1 only to center the pinhole. 1. RM2 columns have pulling-down formula to calculate desired RM2 positions based on the Az/El measurement in the above row: = (target AZ/AL -measured AZ/EL)* overcorrection constant. You may also overwrite the calculation with numbers to calculate in the next row; You may add new rows. Do not change or erase the equations above your rows.
| Target values --------> |  |  |  | 0 | 0 | 808 | 602 |  |  | 2B/AZ | 2A/EL | 2. Apply the calculated RM2 positions to ADM. And apply the rough estimated RM1 values.
2023-02-19 | 20.5205 | 23.804 | 21.8163 | 23.9211 | 0.0076 | -0.0069 | 808.4 | 602.6 |  | 0.5 | 10.6 | -17.9 | 3. Use RM1 only to move pinhole to center. Record centroid values and Az/El values.
|  |  | 21.6928 | 23.8405 | [5 empty] | PETE: Measure LED first |  |  | 4. If need to do another iterations, pull down the RM2 formula to do the next row.
2023-02-24 | 20.4755 | 23.805 | 21.6928 | 23.8405 | -0.0813 | -0.0492 | 807.9 | 601.8 |  |  | 10.6 | -17.9 | *If the Az/El are too large, i.e., > 0.06, cut down the overcorrection numbers, 10.6 and -17.9 by half to 5 and -8. Do a first iteration to see the improvement and then you need a 2nd iteration with the correct overcorrection of 10.6 and -17.9.
|  |  | 20.8121 | 24.7023
2023-02-24 | 20.937 | 23.489 | 21.2549 | 24.2714 | -0.0379 | -0.0235 | 807 | 602.6
2023-02-24 | 21.367 | 23.197 | 20.8343 | 24.6731 | -0.0015 | 0.0053 | 807.7 | 601.8
2023-02-25 | 21.367 | 23.197 | 20.8343 | 24.6731 | -0.0015 | 0.0084 | 809.5 | 602.6
2023-02-25 | 21.39 | 23.301 | 20.9847 | 24.689 | 0.0041 | -0.0086 | 806.6 | 602.6 |  |  | 10.6 | -17.9
2023-02-25 | 21.3475 | 23.201 | 20.8307 | 24.6455 | -0.0053 | 0.0067 | 808.5 | 601.7
2023-02-25 | 21.3915 | 23.2385 | 20.8863 | 24.6866 | -0.0028 | 0.0032 | 807.5 | 601.6
2023-02-25 |  |  | 20.9436 | 24.7163
2023-02-27 | 21.3915 | 23.2385 | 20.8863 | 24.6866 | -0.0278 | 0.0244 | 825.9 | 579.6
2023-02-27 | 21.5371 | 23.323 | 21.3231 | 24.9813 | 0.0117 | -0.0097 | 807.7 | 602.5 | 20.8863 | 24.6866
2023-02-27 | 21.5601 | 23.403 | 21.1494 | 24.8573 | 0.0169 | -0.0145 | 808 | 602.6
2023-02-27 | 21.3731 | 23.2258 | 20.8899 | 24.6781 | -0.0015 | 0.0028 | 806.6 | 603.8
2023-02-27 |  |  | 20.94 | 24.694 |  |  | 807.8 | 601.6
2023-03-15 |  |  | 20.8863 | 24.6866 | 0.0833 | 0.0508 | 807.6 | 602.6
2023-03-15 |  |  | 21.3384 | 24.2451 | 0.0399 | 0.0247 | 807.5 | 601.5
2023-03-15 |  |  | 21.7805 | 23.8222 | -0.0001 | -0.0034 | 807.6 | 602.5
2023-03-31 | 20.441 | 23.892 | 21.781 | 23.822 | 0.0322 | 0.0147 | 806.755 | 601.684
2023-03-31 | Didn't record after pinhole centered |  | 22.0441 | 23.4807 | ~0.4, or image too bad to measure |  | 807.624 | 601.7
2023-03-31 | 20.441 | 23.892 | 21.7805 | 23.8222 | 0.0093 | 0.0025 | 794 | 582
2023-03-31 | 20.441 | 23.892 | 21.781 | 23.822 | 0.0322 | 0.0133 | 807.4 | 601.8 | Restart at row 23, hoping something avoidable error happed in row 24
2023-03-31 | 20.0655 | 243.074 | 22.0191 | 23.4807 | 0.0055 | 0.006 | 812.8 | 606.4
2023-04-03 | 20.441 | 23.892 |  | 23.822 | 0.1045 | -0.0006 | 808.6 | 666.7 |  |  | 10.6 | -17.9
2023-04-03 | 20.494 | 23.892 | 21.781 | 23.822 | 0.0274 | 0.0001 | 808.6 | 601.8 |  |  | 5.3 | -8.95
2023-04-03 | 20.3385 | 23.8945 | 21.7828 | 23.5316 | 0.0135 | -0.0006 | 807.8 | 601.9
2023-04-03 | 20.261 | 23.8925 | 21.7721 | 23.3885 | 0.0065 | -0.0006 | 807.8 | 601.7 |  |  | sticking with 5.3, -8.95 overcorrection seems to work much better even when the values are small
2023-04-03 | 20.223 | 23.8925 | 21.7613 | 23.3196 | 0.002 | 0.0008 | 807.6 | 601.8
2023-04-06 | 20.223 | 23.893 | 21.771 | 23.571 | -0.0747 | 0.0091
2023-04-06 | 20.6015 | 23.9425 | 21.9339 | 24.3628 | 0.0465 | 0.0098 | 807 | 601.9
2023-04-06 | 20.3354 | 24.0075 | 22.1093 | 23.8699 | 0.0235 | 0.0053 | 808.7 | 601.8
2023-04-06 | 20.203 | 24.0431 | 22.2042 | 23.6208 | 0.0083 | 0.0046 | 807.7 | 601.8
2023-04-06 | 20.1561 | 24.0712 | 22.2865 | 23.5328 | 0.0041 | 0.0004 | 807.6 | 602.6
2023-04-10 | 20.1561 | 24.0712 | 22.0288 | 23.5519 | 0.0801 | 0.0001 |  |  |  |  | 21.2 | -35.8
2023-04-10 | 20.755 | 24.0522 | 22.0297 | 24.0474 | 0.0513 | -0.0176 | 808 | 602
2023-04-10 | 20.457 | 23.9522 | 21.7147 | 23.5036 | 0.026 | -0.009 | 807.6 | 601.9 |  |  | RM2B | RM2A
2023-04-10 | 20.3054 | 23.9016 | 21.5536 | 23.228 | 0.0128 | -0.0044 | 806.8 | 601.9 |  |  | 0.17702 | -0.51552
2023-04-10 | 20.2274 | 23.8756 | 21.4748 | 23.0923 | 0.0048 | -0.0013 | 806.8 | 601 |  |  | 23.7469 | 21.2368
2023-05-23 | 20.2274 | 23.8756 | 21.7523 | 23.5699 | -0.0167 | -0.0288 | 808.439 | 601.672
|  |  | 21.2368 | 23.7469 | 0.0003 | 0.0049
2023-05-24 | 20.312 | 23.7254 | 21.5236 | 23.6494 | -0.0022 | 0.0028 | 807.48 | 601.85
|  |  | 21.5737 | 23.6727
2023-05-25 | 20.3048 | 23.7255 | 21.5236 | 23.6494 | 0.0107 | 0.0049 | 807.598 | 601.839 |  |  |  |  | POST SHIM NULL
|  |  | 21.6113 | 23.536
2023-05-25 | 20.5205 | 23.804 | 21.8163 | 23.9211 | 0.0076 | -0.0069 | 807.86 | 601.75
|  |  | 21.6928 | 23.8405
2023-05-26 | 20.2408 | 23.7535 | 21.5675 | 23.5927 | -0.0654 | -0.06 | 807.9 | 602.6
|  |  | 20.4935 | 24.2859
2023-05-26 | 20.8118 | 23.2155 | 20.4935 | 24.2859 | -0.0334 | -0.0263 | 807.8 | 601.7
|  |  | 20.0227 | 24.64
2023-05-26 | 20.8118 | 23.2155 | 20.0227 | 24.64 | -0.0188 | -0.0097 | 807.7 | 601.74
|  |  | 19.8491 | 24.8393
2023-05-26 | 20.9817 | 23.1325 | 19.8491 | 24.8393 | -0.0049 | -0.0013 | 807.7 | 601.9
|  |  | 19.8258 | 24.8912
2023-05-31 | 20.9858 | 23.1351 | 20.6689 | 24.2679 | -0.00136 | 0.0053 | 814.7 | 592.9
|  |  | 20.7638 | 24.2823
2023-05-31 | 20.9915 | 23.1675 | 20.7163 | 24.2751 | -0.0147 | 0.0035 | 814.6 | 592.66
|  |  | 20.779 | 24.4309
2023-05-31 | 21.0759 | 23.1874 | 20.7476 | 24.353 | -0.0049 | 0.0008 | 814.6 | 592.96
|  |  | 20.7619 | 24.4049
2023-06-01 | [6 empty] | 808 | 602
2023-06-01 | 21.104 | 23.1926 | 20.7548 | 24.379 | -0.0063 | 0.006 | 820.78 | 588.596 | << WRONG X,Y CENTROID TARGET
|  |  | 20.8622 | 24.4458
2023-06-01 | 21.1718 | 23.2484 | 20.8515 | 24.4391 | 0.0013 | -0.0031 | 814.05 | 593.88 | << WRONG X,Y CENTROID TARGET
|  |  | 20.796 | 24.4253
2023-06-01 | 21.54 | 23.2098 | 20.8016 | 24.4267 | 0.0069 | -0.0058 | 814.6 | 594.6 | << WRONG X,Y CENTROID TARGET
|  |  | 20.6978 | 24.3536
2023-06-01 | 21.1132 | 23.1748 | 20.7497 | 24.3901 | 0.009 | -0.0076 | 807.68 | 601.9 |  |  |  |  | CONSIDER LED X,Y TARGET VALUES OF 804,606 FOR RM1 DISABE RELAXATION
|  |  | 20.6137 | 24.2947
2023-06-01 | 20.2972 | 23.7116 | 21.5236 | 23.6494 | 0.0173 | 0.008 | 807.85 | 601.81
|  |  | 21.6668 | 23.466 |  |  |  |  | dPixX | dPixY
2023-06-01 | 20.302 | 23.7142 | 21.5236 | 23.6494 | 0.0103 | 0.0049 | 812 | 609 | 4 | 7
| 20.1685 | 23.7521 | 21.6113 | 23.5402
2023-06-01 | 20.184 | 23.7764 | 21.6113 | 23.5402 | 0 | 0 | 812 | 609
| 20.164 | 23.7564 | 21.6113 | 23.5402
2023-06-01 | 20.1892 | 23.7984 | 21.6113 | 23.5402 | 0.002 | -0.0024 | 812 | 609 |  | 0.001 | -0.0012
| 20.1472 | 23.75 | 21.5683 | 23.519
2023-06-01 | 20.1683 | 23.7452 | 21.5543 | 23.5222 | 0.0024 | -0.0027 | 812 | 609
| 20.1218 | 23.6933 | 21.506 | 23.4968
| 20.1018 | 23.6733
2023-06-05 | 20.0818 | 23.6533 | 21.5619 | 23.6801 | -0.0011 | 0.0032 | 807.67 | 601.56
2023-06-13 | 20.348 | 23.624 | 21.53 | 23.67 | -0.0452 | 0.016 | 808.47 | 603.16
| 20.8263 | 23.793 | 21.8164 | 24.1491 |  |  | 807.362 | 602.373
| 20.8685 | 23.8145 | 21.8164 | 24.149 | 0.0103 | 0.0004 | 807.315 | 601.027
| 20.735 | 23.7992 | 21.8236 | 24.0398
2023-06-16 | 20.7545 | 23.8135 | 21.8236 | 24.0398 | -0.0108 | 0.0001
| 20.8536 | 23.7947 | 21.8254 | 24.1543 |  |  | 807.624 | 601.66
| 20.8336 | 23.7747 | 21.8254 | 24.1543
| 20.8136 | 23.7547 | 21.8254 | 24.1543
| 20.7936 | 23.7347
| 20.7736 | 23.7147
| 20.7536 | 23.6947
2023-07-24 | 20.8835 | 23.792 | 21.83 | 24.15 | -0.0004 | -0.0062 | 808 | 602
2023-07-24 | 20.8885 | 23.757 | 21.719 | 24.1542 | -0.0028 | 0.0001 | 808 | 602
2023-07-24 | 20.892 | 23.759 | 21.7754 | 24.167 | -0.0063 | 0.0032
2023-07-24 | 20.9765 | 23.796 | 21.8327 | 24.2337 | 0.0048 | -0.0017
2023-07-24 | 20.9236 | 23.7765 | 21.8023 | 24.1829 | -0.0022 | 0.0018
2023-07-24 | 20.9347 | 23.7865 | 21.8345 | 24.2062 | 0.002 | -0.0017
2023-07-24 | 20.9241 | 23.7769 | 21.804 | 24.185 | 0.0006 | -0.0006
|  |  | 21.7933 | 24.1786
|  |  | 21.7933 | 24.1786
2023-08-07 |  |  | 22.2773 | 23.7641 | 0.0027 | 0.0032
|  |  | 22.3346 | 23.7355 | -0.0004 | 0.0001
|  |  | 22.3364 | 23.7397 | 0.0003 | 0.0001
2023-08-12 |  |  | 22.0705 | 23.8532 | -0.0094 | -0.0103
|  |  | 21.8861 | 23.9528
2023-09-07 |  |  | 21.4035 | 24.5943 | 0.1006 | 0.0077
|  |  | 21.5413 | 23.5279
2023-09-08 |  |  | 21.443 | 23.858 | 0.0076 | 0.0285
|  |  | 21.9532 | 23.7774

## RxRy Tuner
Date | Rx (deg) | Ry (deg) | AC Az (deg) | AC EL (deg) | Target Az(deg) | Target El(deg) | Angle between Az and Rx (deg) | (rad.) | 1. RxRy columns have pulling-down formula to calculate desired positions from the Az/El measurement & target in the above row. The equations use the rotational transformation of a rotational angle. Do not arbitrarily change the angle. You may add new rows. Do not change or erase the equations above your rows.
| 0.0907 | 0.0813 | 0.0228 | 0.0501 | -0.0001 | 0.0042 | 34.5 | 0.60214 | 2. Apply the calculated RX Ry to 5DOF.
Go here -> | 0.045829 | 0.056443 |  |  |  |  | Do not change thses! |  | 3. Record the Az/El measurement.
4. If need to do another iteration, fill in the target Az/El, the next row calcutates the new Rx Ry.

## X Y Tuner
Date | dX (mm) | dY (mm) | Px (pix) | Py(pix) | Target Px (pix) | Target Py (pix) | Angle between Az and Rx (deg) | (rad.) | Scale (pix/mm) | 1. dX, dY columns have pulling-down formula to calculate desired positions from the Px/Py measurement & target in the above row. The equations use the rotational transformation of a rotational angle. Do not arbitrarily change the angle or scale. You may add new rows. Do not change or erase the equations above your rows.
| 0 | 0 | 808.8 | 599.6 | 808.8 | 599.9 | 34.5 | 0.60214 | 20.9 | 2. Apply the calculated increments dX, dY to 5DOF.
Deltas -> | -0.0081302 | -0.01183 |  |  |  |  | Do not change thsese! |  |  | 3. Record the Px/Py measurements.
| 0 | 0 | 813.282 | 597.999 | 814.584 | 597.53 |  |  |  | 4. If need to do another iteration, fill in the target Px/Py, the next row calcutates the new dX dY.
| 0.064051 | -0.016792
| 0 | 0 | 824.152 | 590.625 | 814.584 | 597.53
| -0.56442 | -0.012977
| 0 | 0 | 818.574 | 608.031 | 812.681 | 597.662
| 0.048636 | 0.56857 |  |  | 812.201 | 598.029
| 15.8195 | -45.5927
| 0 | 0 | 819 | 565 | 808 | 602
| -1.4365 | -1.1609

## Position Log - GSFC
Date | Time |  | RM1 |  | RM2 |  | RM3 |  | 5DOF x(mm) | 5DOF y(mm) | 5DOF z(mm) | 5DOF Rx(deg) | 5DOF Ry(deg) | AC Focus | Cam exp (ms) | Lidar Sig | Range(m) | .fits | .bmp | AC AZ | AC EL | X centroid | Y centroid | Radius | Notes | Time(EST) | ADMBNCHT | ADMRADT | ADMBOXT | ADMAIRT | ADMCAMT | ADMWINT | ADMLEICA
|  |  | A | B | A | B | A | B | [9 empty] | Infinity
3:385/14/2022 |  | MATF Retro1 | 0 | 0 | 0 | 0 | -3700 | 5000
RM3 rezero
2022-09-16 |  | MATF Retro1 | 0 | 0 | 0 | 0 | -600 | 200 | [7 empty] | 31 | 4.0315
2022-09-16 |  | SMPA Retro2 | 0 | 0 | 0 | 0 | 7000 | -8100 | [7 empty] | 31 | 4.7625
2022-09-16 |  | Focus at mirror | -2000 |  | -500 | -700 | [7 empty] | 0,202
RM, focus rezero
2022-09-19 |  | focus Retro2 |  |  | 7200 | 3100 | [7 empty] | 1, 19 |  | 40 | 4.7625
2022-09-19 |  | Focus at mirror, center illumination to code bit | -2000 |  | -500 | -700 | [7 empty] | 1,77
2022-09-20 |  | Retro 1 range |  |  |  |  | -4300 | 5200 | [7 empty] | 31 | 4.03
2022-09-20 |  | SMPA Retro2 |  |  |  |  | 3300 | -3100 | [7 empty] | 31 | 4.7625
2022-09-20 |  | SMPA Retro2 |  |  |  |  | 3300 | -3100 | [5 empty] | 1,69
2022-09-20 |  | center illumination from north east to mirror center | 20000 | 4000 | 4000 | 20000 | 3300 | -3100 | [5 empty] | 1,68
2022-09-20
2022-09-20 |  | Good infinity with theodolite looking back | [11 empty] | 0,20
2022-09-20 |  | focus Retro1 |  |  | -9000 | -5000 | [7 empty] | 1,73
2022-09-20 |  | Inifinity Retro1, faint but fair code bit image |  |  | -9000 | -5000 | [7 empty] | 0,25
2022-09-20 |  | inifinity Mirror, faint and blurred code bit pattern | 20000 | 4000 | 4000 | 20000 | [7 empty] | 0,20
2022-09-20 |  | inifinity Mirror, faint, pattern washed out | 0 | 0 | 0 | 0 | [7 empty] | 0,20
2022-09-20 |  | Inifinity Retro1, washed out | 20000 | 4000 | -5000 | 15000 | [7 empty] | 0,25
2022-09-20 |  | lost mirror position
2022-09-21 |  | Retro2 |  |  | -1000 | -2000 | [9 empty] | 35 | 4.76
2022-09-21 |  | Mirror |  |  | -8000 | -5000
2022-09-21 |  | Rezero M2 to about on 9/20 (green line)
2022-09-21 |  | infininity Retro1, better image |  |  | -10400 | -6000 | [7 empty] | 0,28
2022-09-21 |  | Retro2 |  |  | 8000 | 3000
2022-09-21 |  | Theodolite centers laser to code-bits |  |  |  |  | 750 | 500
2022-09-21 |  | Rezero M3 to green line
2022-09-21 |  | Center illumination &pinhole to CCD | 23000 | 1500 | 2000 | 21000
2022-09-22 |  | Moving Salsa mirror torward center, Lidar beam started to clip middle way; moving RM1 + in A&B gave more room for Salsa turning
2022-09-22 |  |  | 28000 | 3500
2022-09-22 |  | lost mirror position, probably rezeroed nearby green line
2022-09-22 |  | checking inifinity, bits image improved | 1000 | 6000
2022-09-22 |  |  | 1000 | 12000 | -1000
2022-09-22 |  |  | 1000 | 0 | -18000 | 1000 | [11 empty] | 51711.fits
2022-09-22 |  |  | 1000 | -1500 | -20000 | 1000 | [11 empty] | 53258
2022-09-22 |  |  | 1000 | -4000 | -24000 | 1000 | [11 empty] | 54031
2022-09-22 |  |  | 1000 | -6900 | -28000 | 1000 | [11 empty] | 54559
2022-09-22 |  |  | 1000 | -9500 | -32000 | 1000 | [11 empty] | 55020
2022-09-22 |  |  | 1500 | -12000 | -35500 | 1000 | [11 empty] | 60423
2022-09-23 |  | Jeo had to laterally shift the theodolite to the center of ADM | 1500 | -12000 | -35500 | 1000 | [7 empty] | 0,20 |  |  |  | 21526
2022-09-23 |  | tapped Salsa slightly, the laser move from NW to center at mirror focus and move from SW closer to center | 1500 | -12000 | -36000 | 1000 | [7 empty] | 0,20 |  |  |  | 33307
|  | Retro1 | -18500 | -16000 | -40000 | -16500 | [7 empty] | 0,20 |  |  |  | 50148
2022-09-23 |  | Save image with laser on | 1500 | -12000 | -36000 | 1000 | [7 empty] | 1,71
2022-09-23 |  | Save image with laser on | 1500 | -12000 | -36000 | 1000 | [7 empty] | 0,20 |  |  |  | 54140
|  |  | -18500 | -16000 | -40800 | -17000 | [9 empty] | 17 | 4.02
2022-09-23 |  | Save image with IR laser on | 1500 | -12000 | -36000 | 1000 | [7 empty] | 1,62
2022-09-23 |  | Save image with IR laser on | 1500 | -12000 | -36000 | 1000 | [7 empty] | 0,20 |  |  |  | 65018
ADM overhaul replacing all 6 RM mirror actuators, and the LEDs
2022-12-14 |  | Initial good RM positions | 21 | 22 | 19.02 | 22 | 22 | 22
2022-12-14 |  |  | 21 | 22 | 19.02 | 22 | 22 | 25
2022-12-14 |  |  | 21 | 22 | 19.02 | 22 | 22 | 19
2022-12-14 |  |  | 21 | 22 | 19.02 | 22 | 22 | 22
2022-12-14 |  |  | 21 | 22 | 19.02 | 22 | 24.8 | 22
2022-12-14 |  |  | 21 | 22 | 19.02 | 22 | 24.8 | 24
2022-12-14 |  | laser centers on autocollimator | 21 | 22 | 19.02 | 22 | 24.8 | 23.5
2022-12-16 |  |  | 21 | 22 | 19.02 | 22 | 24.8 | 23.5 | [5 empty] | 0,28
2022-12-16 |  |  | 21 | 22 | 19.02 | 22 | 24.8 | 23.5 | [5 empty] | 0,0 |  |  |  | 12209
2022-12-16 |  |  | 21.5 | 22 | 19.02 | 22 | 24.8 | 23.5 | [5 empty] | 0,1 |  |  |  | 13448
2022-12-16 |  |  | 22 | 22 | 19.02 | 22 | 24.8 | 23.5 | [5 empty] | 0,2 |  |  |  | 13807
2022-12-16 |  |  | 22.5 | 22 | 19.02 | 22 | 24.8 | 23.5 | [5 empty] | 0,3 |  |  |  | 14052
2022-12-16 |  |  | 23 | 22 | 19.02 | 22 | 24.8 | 23.5 | [5 empty] | 0,4 |  |  |  | 14323
2022-12-16 |  | rm1a optimal at 22 | 22 | 22 | 19.02 | 22 | 24.8 | 23.5 | [5 empty] | 0,5 |  |  |  | 14648
2022-12-16 |  | worse | 22 | 21.5 | 19.02 | 22 | 24.8 | 23.5 | [9 empty] | 15039
2022-12-16 |  |  | 22 | 22.5 | 19.02 | 22 | 24.8 | 23.5 | [9 empty] | 15250
2022-12-16 |  |  | 22 | 23 | 19.02 | 22 | 24.8 | 23.5 | [9 empty] | 15540
2022-12-16 |  |  | 22 | 23.5 | 19.02 | 22 | 24.8 | 23.5 | [9 empty] | 15901
2022-12-16 |  |  | 22 | 24 | 19.02 | 22 | 24.8 | 23.5 | [9 empty] | 20222
2022-12-16 |  |  | 22 | 23.5 | 19.02 | 22 | 24.8 | 23.5 | [9 empty] | 20633
2022-12-16 |  |  | 22 | 23.5 | 19.5 | 22 | 24.8 | 23.5 | [9 empty] | 20932
2022-12-16 |  |  | 22 | 23.5 | 20 | 22 | 24.8 | 23.5 | [9 empty] | 21116
2022-12-16 |  |  | 22 | 23.5 | 20.5 | 22 | 24.8 | 23.5 | [9 empty] | 21333
2022-12-16 |  |  | 22 | 23.5 | 21 | 22 | 24.8 | 23.5 | [9 empty] | 21707
2022-12-16 |  |  | 22 | 23.5 | 21.5 | 22 | 24.8 | 23.5 | [9 empty] | 21914
2022-12-16 |  | rm2b might have missed move | 22 | 23.5 | 20.5 | 22.5 | 24.8 | 23.5 | [9 empty] | 22726
2022-12-16 |  |  | 22 | 23.5 | 20.5 | 22.5 | 24.8 | 23.5 | [9 empty] | 22941
2022-12-16 |  |  | 22 | 23.5 | 20.5 | 22.5 | 24.8 | 23.5 | [5 empty] | 1,45 |  |  |  | 42537
2022-12-16 | [13 empty] | 2100 |  |  |  | 45315
2022-12-16 |  | smr image | [11 empty] | 2200 |  |  |  | 45631
2022-12-16 | [13 empty] | 2150 |  |  |  | 45821
2022-12-16 | [13 empty] | 2,75 |  |  |  | 45957
2022-12-16 | [13 empty] | 2100 |  |  |  | 50107
2022-12-16 |  |  |  | 23.7 | [13 empty] | 50533
2022-12-16 |  |  | 22.2 | [14 empty] | 51044
2022-12-16 |  |  | 21.8 | [14 empty] | 51605
2022-12-16 |  |  | 21.6 | [14 empty] | 51715
2022-12-16 |  |  | 21.4 | [14 empty] | 51815
2022-12-16 |  |  | 21.2 | [14 empty] | 51930
2022-12-16 |  |  | 21 | [14 empty] | 52056
2022-12-16 |  |  |  | 23.9 | [13 empty] | 52311
2022-12-16 |  |  |  | 23.5 | [13 empty] | 52532
2022-12-16 |  |  |  | 23.3 | [13 empty] | 52647
2022-12-16 | [13 empty] | 0,20 |  |  |  | 52822
2022-12-16 | [6 empty] | 24.4 | 23.2 | [5 empty] | 2100 |  |  |  | 54135
2022-12-16 | [13 empty] | 2,0 |  |  |  | 54506
2022-12-16 |  |  | 21 | 23.3 | 20.5 | 22.5 | 24.4 | 23.2 | [5 empty] | 2100 |  |  |  | 70420
2022-12-17 |  |  | 21 | 23.3 | 21.1 | 23.95 | [7 empty] | 0,0 |  |  |  | 20909
2022-12-17 |  |  | 21 | 23.3 | 21.1 | 24.7 | [7 empty] | 0'0 |  |  |  | 21621
2022-12-17 |  |  | 21 | 23.3 | 21.1 | 24.7 | [7 empty] | 2100 |  |  |  | 22629
2022-12-17 |  | Theodolite aligned | 21 | 23.3 | 21.1 | 24.7 | [7 empty] | 0,20 |  |  |  | 23215
2022-12-17 |  | flat mirror image | 21.221 | 22.897 | 20.389 | 24.884 | [7 empty] | 0,20 |  |  |  | 71359
2022-12-18 |  |  | 21.221 | 22.897 | 20.389 | 24.884 | 24.4 | 23.2 | [5 empty] | 2100
2022-12-18 |  |  | 21.221 | 22.897 | 20.389 | 24.884 | 24.4 | 23.2 | [5 empty] | 2118
2022-12-18 |  |  | 21.221 | 22.897 | 20.389 | 24.884 | 24.4 | 23.2 | [5 empty] | 0,24 |  |  |  | 92852
2022-12-18 |  |  | 21.321 | 23 | 20.5 | 24.9 | 24.4 | 23.2 | [5 empty] | 2118 |  |  |  | 13149
2022-12-18 |  |  | 21.321 | 23 | 20.5 | 24.9 | 24.4 | 23.2 | [5 empty] | 2118 | [5 empty] | -0.0108 | 0.1237
[5 empty] | 20.51 | [14 empty] | -0.0091 | 0.1119
[5 empty] | 20.56 | [14 empty] | -0.0094 | 0.0511
2022-12-18 |  | autocollimation on close flat | 21.321 | 23 | 20.61 | 24.9 | 24.4 | 23.2 | [5 empty] | 0,24 | 5000 |  |  | 40344 |  | -0.0098 | -0.0131
2022-12-18 |  | laser on close flat | 21.321 | 23 | 20.61 | 24.9 | 24.4 | 23.45 | [5 empty] | 0,24 | 250 |  |  | 40312
2022-12-18 |  |  | 21.071 | 22.8 | 20.51 | 24.85 | [11 empty] | 65041
2022-12-18 |  | center pinhole to ccd from southwest | 21.071 | 22.8 | 20.25 | 24.65 | [7 empty] | 1,70 |  |  |  | 65646
2022-12-18 |  |  | 21.121 |  |  | 24.7 | [7 empty] | 0,24 |  |  |  | 65913
2022-12-18 |  |  | 21.171 |  |  | 24.75
2022-12-18 |  |  | 21.271 |  |  | 24.83
2022-12-18 |  | keep pinhole centered, (1a/2b)(1b/2ax1.5) | 21.671 | 23.1 | 20.7 | 25.23 | [7 empty] | 0,24 |  |  |  | 75319
2022-12-18 |  | Retro | 21.671 |  | 20.7 | 25.23 | [7 empty] | 0,25 |  |  |  | 93544
2022-12-18 |  | flat mirror | 21.671 | 23 | 20.7 | 25.23 | [7 empty] | 0,25 |  |  |  | 95352
2022-12-18 |  | Lidar signal improved to 3 | 21.671 | 23 | 20.7 | 25.23 | 24.345 | 24.445
2023-01-06 |  | Centered laser , Retro at ADM window | 21.671 | 23 | 20.7 | 25.23 | 23.2 | 23.8 | [5 empty] | 2115 | 500 |  |  | 62847
2023-01-06 |  | center theodolite | 21.65 | 23 | 20.6 | 25.23 | [7 empty] | 0,25
2023-01-06 |  | Theodolite aligned with Laser | 21.65 | 23 | 20.6 | 25.23 | 24.4 | 23.45
2023-01-06 |  | flat mirror aligned with theodolite | 21.65 | 23 | 20.6 | 25.23 | 24.4 | 23.45 | [5 empty] | 2115 | 5000 |  |  | 80437
2023-01-06 |  | flat mirror at infinite focus | 21.65 | 23 | 20.6 | 25.23 | 24.4 | 23.45 | [5 empty] | 0,25 | 5000 |  |  | 80832
2023-01-06 |  | Initial cat's eye sWFI mirror | 21.65 | 23 | 20.6 | 25.23 | 24.4 | 23.45 | [5 empty] | 1,70 | 5000 |  |  | 83154
2023-01-06 |  | sWFI mirror at infinity | 21.65 | 23 | 20.6 | 25.23 | 24.4 | 23.45 | [5 empty] | 0,25 | 5000 |  |  | 83504
2023-01-06 |  |  | 21.65 | 23 | 20.6 | 25.23 | 24.4 | 23.45 | [5 empty] | 0,28 | 5000 |  |  | 83655
2023-01-06 |  |  | 21.65 | 23 | 20.6 | 25.23 | 24.4 | 23.45 | [5 empty] | 0,31 | 5000 |  |  | 83745
2023-01-06 |  |  | 21.65 | 23 | 20.6 | 25.23 | 24.4 | 23.45 | [5 empty] | 0,33 | 5000 |  |  | 84206
2023-01-09 |  | sMATF mirror AC image at infinity | 21.65 | 23 | 20.6 | 25.23 | 24.4 | 23.45 | [5 empty] | 0,31 | 5000 |  |  | 4724
2023-01-09 |  | sMATF mirror at infinity, ADM under some human perturbation | 21.65 | 23 | 20.6 | 25.23 | 24.4 | 23.45 | [5 empty] | 0,32 | 5000 |  |  |  | 48_08 | -0.0063 | -0.0048
2023-01-09 |  | sMATF mirror at infinity, ADM under some human perturbation | 21.65 | 23 | 20.6 | 25.23 | 24.4 | 23.45 | [5 empty] | 0,33 | 5000 |  |  | 5745
2023-01-09 |  |  | 21.65 | 23 | 20.6 | 25.23 | 24.4 | 23.45 | [5 empty] | 0,34 | 5000 |  |  |  | 02_20 | -0.0046 | -0.0062
2023-01-09 |  | ADM under some human perturbation | 21.65 | 23 | 20.6 | 25.23 | 24.4 | 23.45 | [5 empty] | 0,35 | 1000 |  |  |  | 06_32 | -0.006 | -0.0041
2023-01-09 |  | LED13 (sMATF Target) on, measure mirror center | 21.65 | 23 | 20.6 | 25.23 | 24.4 | 23.45 | [5 empty] | 1,71 | [7 empty] | 746.554 | 602.998
2023-01-09 |  | LED13 (sMATF Target) on, measure mirror center | 21.65 | 23 | 20.6 | 25.23 | 24.4 | 23.45 | [9 empty] | 12552 |  |  |  | 745.872 | 602.221
2023-01-09 |  | LED14 (sMATF Target) on, measure mirror center | 21.65 | 23 | 20.6 | 25.23 | 24.4 | 23.45 | [13 empty] | 746.55 | 603.45
2023-01-09 |  | sMATF Flat mirror AC image, cat's eye | 21.65 | 23 | 20.6 | 25.23 | 24.4 | 23.45 | [5 empty] | 2117 | 1000 |  |  | 42547
2023-01-09 |  | sMATF Flat mirror AC image, infinity | 21.65 | 23 | 20.6 | 25.23 | 24.4 | 23.45 | [5 empty] | 0,24 | 1000 |  |  | 42917
2023-01-09 |  | Range, saved MATFrangeJan9.txt | 21.65 | 23 | 20.6 | 25.23 | ? | ? | [7 empty] | 15 | 4.0318
2023-01-09 |  | Tool flat mirror at ADM, cat's eye | 21.65 | 23 | 20.6 | 25.23 | [7 empty] | 1122 | 1000 |  |  | 53827
2023-01-09 |  | Tool flat mirror at ADM, infinity | 21.65 | 23 | 20.6 | 25.23 | [7 empty] | 0,21 | 1000 |  |  | 54259
2023-01-09 |  |  | 21.65 | 23 | 20.6 | 25.23 | [7 empty] | 1122 |  |  |  |  | 44_49 | -0.0629 | -0.1163
2023-01-09 |  |  | 21.65 | 23 | 20.6 | 25.23 | [7 empty] | 1122 |  |  |  |  | 44_49 | -0.0633 | -0.1169
2023-01-09 |  | At nominal 5DOF position, measure PAT2 Primary LED centroid | 21.65 | 23 | 20.6 | 25.23 | [7 empty] | 1122 | [7 empty] | 770.83 | 155.064
2023-01-09 |  | At nominal 5DOF position, measure PAT2 Redundant LED centroid | 21.65 | 23 | 20.6 | 25.23 | [7 empty] | 1122 | [7 empty] | 824.3 | 1115.5
2023-01-09 | [13 empty] | 1122 | [7 empty] | 808 | 602
2023-01-09 |  | PAT2 Primary Retro next to Primary LED, 5DOF moved x=6mm, Y=-26mm from nominal SPA position | [11 empty] | 1122 |  |  | 3.4584
2023-01-10 |  | Re-range PAT2 Primary Retro, assuming RMs parked after prior day's shut down; | [13 empty] | 23 | 3.4583
2023-01-10 |  | Using Lidar IR illumination to take PAT2 primary retro image | 21.16 | 25.34 | [9 empty] | 1158 | 250 |  |  | 4212 |  |  |  | 868 | 536.5
2023-01-10 |  | Move x=105.75,y=101.2, so lidar illuminated PAT2 primary Retro image becomes better; eyeball the apex again | [11 empty] | 1158 | 250 |  |  | 5413 |  |  |  | ? | ?
2023-01-10 |  | Move x=128.493,y=141.205, range PAT2 redundant Retro | [14 empty] | 3.4664
2023-01-10 |  | Take redundant Retro image using Lidar IR beam; eyeball the retro apex off the image | [11 empty] | 1139 |  |  |  | 14706 |  |  |  | 824 | 576
2023-01-10 |  | H2 position, adjust RM2 to center sMATF mirror pinhole to CCD | 21.65 | 23 | 20.49 | 25.225 | 24.4 | 23.45
2023-01-10 |  | Adjust Rm3 to range sMATF retro |  |  |  |  | 24.1469 | 23.75 | [7 empty] | 10 | 4.0317
2023-01-10 |  | Take retro image using Lidar beam |  |  |  |  | 24.1469 | 23.75 | [9 empty] | 42543 |  |  |  | 156 | 1039
2023-01-10 |  | Take AC image including sMATF mirror pinhole and retro/lidar at SW corner |  |  |  |  | 24.1469 | 23.75 | [9 empty] | 43243
2023-01-10 |  | Take AC image including sMATF mirror pinhole and retro/lidar after restoring RM3 positions | 21.65 | 23 | 20.49 | 25.225 | 24.4 | 23.45 | [9 empty] | 43531
2023-01-10 |  | AC image of sMATF mirror Cat's eye | 21.62 | 22.95 |  |  | 24.4 | 23.45 | [5 empty] | 1,71 |  |  |  | 43929
2023-01-10 |  | Ac image of sMATF mirror infinity |  |  |  |  | 24.4 | 23.45 | [5 empty] | 0,27 |  |  |  | 44233
2023-01-10 |  | sMATF AL/EL measurement at Cat's eye |  |  |  |  | 24.4 | 23.45 | [11 empty] | -0.3022 | -0.3312
2023-01-10 |  | sMATF AL/EL measurement at infinity |  |  |  |  | 24.4 | 23.45 | [11 empty] | 0.0503 | -0.0232
2023-01-10 |  | sMATF Mirror at Cat's eye, keep pinhole centered, move illumination from right to center | 21.82 | 23.9 | 21.81 | 25.38 | 24.4 | 23.45 | [5 empty] | 1,71 |  |  |  | 54958
2023-01-10 |  | sMATF Mirror at infinity |  |  |  |  | 24.4 | 23.45 | [5 empty] | 0,15 | [5 empty] | -0.0001 | -0.0367
2023-01-10 |  | sMATF Mirror at infinity |  |  |  |  | 24.4 | 23.45 | [5 empty] | 0,15 |  |  |  | 55313
2023-01-10 |  |  | 21.72 |  |  |  | 24.4 | 23.45 | [11 empty] | 0.1354 | -0.0322
2023-01-10 |  | Fine tuning RMs using the matrix tool | 21.76 | 23.44 | 21.16 | 25.34 | 24.4 | 23.45 | [5 empty] | 1,71 | [7 empty] | 797.2 | 606.7
2023-01-10 |  |  | 21.76 | 23.44 | 21.16 | 25.34 | 24.4 | 23.45 | [5 empty] | 0,25 | [5 empty] | 0.008 | -0.002
2023-01-10 |  | Take sMATF mirror image at infinity | 21.76 | 23.44 | 21.16 | 25.34 | 24.4 | 23.45 | [5 empty] | 0,25 | 1000 |  |  | 63307
2023-01-10 |  | Take sMATF mirror image at Cat's eye | 21.76 | 23.44 | 21.16 | 25.34 | 24.4 | 23.45 | [5 empty] | 1,67 | 1000 |  |  | 63840
2023-01-11 |  | Move 5dof to P2 position |  |  |  |  | 24.4 | 23.45 | 101.5 | 125.7 | 0
2023-01-11 |  | Adjust to Rx=?, Ry=?, AC image of PATB flat mirror; on 1/13, retrived Rx=0.08955, Ry=0.04604 |  |  |  |  | 24.4 | 23.45 | 101.5 | 125.7 | 0 |  |  | 0,30 | 1000 |  |  | 92005 |  | 0.0072 | 0.0004
2023-01-11 |  | PATB Redundant LED, 5DOF moved to x=111.505mm, Y=-148.801mm | [6 empty] | 111.505 | 148.801 | 0 | [10 empty] | 797.7 | 606.1
2023-01-11 |  | 5DOF moved to x=124.405mm, Y=142.401mm to "center" PATB redundant retro | [6 empty] | 124.405 | 142.401 | 0 |  |  | 1,150? |  |  | 3.4668 | 844 |  |  |  | 797.4 | 606.3
2023-01-11 |  | 5DOF moved to x=100.693mm, Y=100.398mm to "center" PATB primary retro | [6 empty] | 100.693 | 100.398 | 0 | [5 empty] | 3.4588 | 2209 |  |  |  | 797.5 | 605.5
2023-01-11 |  | 5DOF moved to x=88.503mm, Y=108.398mm to "center" PATB primary LED | [6 empty] | 88.503 | 108.398 | 0 |  |  | 1,32 |  |  |  | 3130 |  |  |  | 797.8 | 606.7
2023-01-11 |  | Move 5dof to P1 position | [6 empty] | -76.1 | 125.7 | 0
2023-01-11 |  | Adjust to Rx=0.6382, Ry=0.3495, AC image of PATA flat mirror; Data retrival on 1/13 showed that the values should have been Rx=0.06382, Ry=0.03495 | [8 empty] | 0 |  |  | 0,18 |  |  |  | 4719 |  | 0.0062 | -0.0044
2023-01-11 |  | Move 5DOF to "center" PATA primary LED | [6 empty] | -60.597 | 108.399 | 0 |  |  | 1130 | [7 empty] | 797.2 | 606.7
2023-01-11 |  | Move 5DOF to "center" PATA primary retro | [6 empty] | -72.502 | 100.199 | 0 |  |  | 1142 |  |  | 3.4588 |  |  |  |  | 795.5 | 607.5
2023-01-11 |  | Move 5DOF to "center" PATA redundant retro | [6 empty] | ? | 142.101? | 0 |  |  | 1134 |  |  | 3.4668 | 13252 |  |  |  | 796.4 | 606.5
2023-01-11 |  | Move 5DOF to "center" PATA redundant LED | [6 empty] | -83.618 | 148.642 | 0 |  |  | 1120 |  |  |  | 13948 |  |  |  | 797 | 606.7
2023-01-11 |  | Adjust to Rx=0.08766, Ry=0.04135, retake center AC image of PATB flat mirror | [6 empty] | 101.5 | 125.7 | 0 |  |  | 0,16 |  |  |  | 15210 |  | 0.0062 | -0.0048
2023-01-12 |  | Retrieve the I186 and J186 using the 5DOF position log and the time stamp | [6 empty] | -96.291 | 142.101
2023-01-21 |  | Add 10 new 5DOF PATA PATB nominal positions
2023-01-21 |  | Move to PATA Mirror, rx=0.06382,ry=0.03495 | [6 empty] | -76.1 | 125.7 |  |  |  | 0,16 | 1000 |  |  | 14647 |  | 0.0041 | -0.002
2023-01-21 |  | PATA LED primary, targeting (797,607) | [6 empty] | -60.656 | 108.448 | [11 empty] | 796.9 | 607.9
2023-01-21 |  | PATA Retro Primary | [6 empty] | -71.5 | 100.2 |  |  |  | 1144 |  |  | 3.4588 | 25636 |  |  |  | 781 | 620 |  | -4e-05
2023-01-21 |  | PATA Retro Redantant | [6 empty] | -97.508 | 142.6 | [6 empty] | 3.4669 | 31536 |  |  |  | 775 | 613 |  | 8.6e-05
2023-01-21 |  | PATA LED Redundant | [6 empty] | -83.76 | 148.69 | [11 empty] | 797.2 | 607
2023-01-21 |  | Move to PATB Mirror, rx=0.08955,ry=0.04604 | [6 empty] | 101.54 | 125.7 |  |  |  | 0,15 |  |  |  | 41034 |  | 0.0076 | 0.0011
2023-01-21 |  | PATB LED primary | [6 empty] | 88.353 | 108.404
2023-01-21 |  | PATB Retro Primary | [6 empty] | 99.702 | 102.4 |  |  |  | 1128 |  | 6 | 3.4592 | 43535 |  | 759 | 581 |  |  |  | 0.000437
2023-01-21 |  | PATB Retro Redundant | [6 empty] | 123.415 | 144.4 |  |  |  | 1124 |  | 6 | 3.4672 | 45219 | [6 empty] | 0.000471
2023-01-21 |  | PATB LED Redundant | [6 empty] | 111.354 | 140.801 | [11 empty] | 796.1 | 606.7
2023-01-21 |  | Move to H2 position
2023-01-21 |  | Check sMATF mirror center | [11 empty] | 1,65 | 2000 |  |  | 55646 |  |  |  | 808.4 | 605.5
2023-01-21 |  | sMATF Mirror AC image, cat's eye | [11 empty] | 1,65 | 2000 |  |  | 55749
2023-01-21 |  | sMATF AC image, infinity focus | [15 empty] | 63535 |  | 0.008 | 0.022
2023-01-21 |  | Adjust ADM rm3 to sMATF Retro |  |  |  |  | 24.15 | 23.75 | [7 empty] | 17 | 4.0322 | 71249 |  |  |  | 131 | 1058 |  | 0.000499
2023-01-21 |  | PATA LED Redundant, tuning 5DOF x&Y to reach new target centroid (808,606, the sMATF reference) | [6 empty] | -83.37 | 148.34 |  |  |  | 1130 | [7 empty] | 808.1 | 606.5
2023-01-21 |  | PATA LED primary | [6 empty] | -60.3 | 108.15 |  |  |  | 1122 | [7 empty] | 808.9 | 605
2023-01-21 |  | PATB LED primary | [6 empty] | 88.75 | 108.15 |  |  |  | 1122 | [7 empty] | 808 | 605.7
2023-01-21 |  | PATB LED Redundant | [6 empty] | 111.7 | 148.5 |  |  |  | 1119 | [7 empty] | 807.7 | 606
2023-01-22 |  | PATB LED Redundant, checking repeatability overnight and rack on/off | [19 empty] | 807 | 607.9
Date |  |  | RM1 |  | RM2 |  | RM3 |  | 5DOFx | 5DOFy | 5DOFz |  |  | AC Focus | Cam exp(ms) | Lidar Sig | Range(m) | .fits | .bmp | AC AZ | AC EL | X centroid | Y centroid |  | Range repeatability (m)
2023-01-23 |  | move from pr 1 H2 position | 21.76 | 23.44 | 21.16 | 25.34 | 24.15 | 23.75 | [11 empty] | 0.026 | 0.067 | 840 | 614
2023-01-23 |  |  | 21.75 | 23.32 | 21.04 | 25.31 | [7 empty] | 0,21 |  |  |  | 1625 |  | 0.004 | 0.011
2023-01-23 |  | sMATF mirror crosshair | [11 empty] | 1,73 |  |  |  | 2250 |  |  |  | 785 | 597
2023-01-23 |  | sMATF Retro | [12 empty] | 500 | 7~ | 4.0322 | 4930 |  |  |  | 134 | 1036 |  | -1.4e-05
2023-01-23 |  | PATAB Mirror | [6 empty] | 101.5 | 125.7 |  |  |  | 0,24 |  |  |  | 24721 |  | -0.004 | -0.014
2023-01-23 |  | PATAB LED | [6 empty] | 89.4 | 107.3 |  |  |  | 1141 | [7 empty] | 806.8 | 606.6 |  | Crosshair ref?
2023-01-24 |  | Post PR2 move to H2 and set RMs | 21.75 | 23.32 | 21.04 | 25.31 | 24.15 | 23.75 | [11 empty] | -0.0261 | -0.043 | 738.8 | 572.7
|  | range to sMATF retro | 21.75 | 23.32 | 21.04 | 25.31 | 24.1 | 23.74 | [16 empty] | visible return in bottom left of camera
|  | tried to use tweak tool by going straight to the last iteration and it did not work. Stopping for shift change | [22 empty] | now AC working
Date |  |  | RM1 |  | RM2 |  | RM3 |  | 5DOFx | 5DOFy | 5DOFz |  |  | AC Focus | Cam exp(ms) | Lidar Sig | Range(m) | .fits | .bmp | AC AZ | AC EL | X centroid | Y centroid |  | Range repeatability (m)
2023-01-25 |  | sMATF mirror AC B227 | 21.75 | 23.32 | 21.04 | 25.31 | 24.15 | 23.75 | [11 empty] | -0.03 | 0.016 | 787 | 571
2023-01-25 |  | Use Evan's Matrix to tune | 21.68 | 21.66 | 21.53 | 25.27 | 24.15 | 23.75 | [9 empty] | 423 |  | -0.0004 | -0.005
2023-01-25 |  | Range sMATF retro | 21.68 | 21.66 | 21.53 | 25.27 | 24.15 | 23.75 | [7 empty] | 9 | 4.0321 | 329 | [6 empty] | -7e-05
2023-01-25 |  | sMATF mirror pinhole | 21.68 | 21.66 | 21.53 | 25.27 | 24.15 | 23.75 | [9 empty] | 136 |  |  |  | 796.8 | 599
2023-01-25 |  | PATB mirror AC rx=0.089, ry=0.046 | 21.68 | 21.66 | 21.53 | 25.27 | 24.15 | 23.75 | [11 empty] | -0.007 | -0.03
2023-01-25 |  | PATB mirror AC rx=0.106, ry=0.061, update all 5 nominals with new Rx, Ry | 21.68 | 21.66 | 21.53 | 25.27 | 24.15 | 23.75 | 101.5 | 125.7 | [7 empty] | 4357 |  | -0.0011 | -0.0107
2023-01-25 |  | PATB primary LED | 21.68 | 21.66 | 21.53 | 25.27 | 24.15 | 23.75 | 88.69 | 108.5 |  |  |  | 1131 |  |  |  | 5730 |  |  |  | 796.8 | 598.6
2023-01-25 |  | PATB primary Retro | 21.68 | 21.66 | 21.53 | 25.27 | 24.15 | 23.75 | 69.2 | 102.9 |  |  |  | 1140 |  |  | 3.4596 | 12238 | [6 empty] | 0.000368
2023-01-25 |  | Apply offsets to all PATA PATB retro nominal table with x:-31.5 and y:+2.5 on ASIST | 21.68 | 21.66 | 21.53 | 25.27 | 24.15 | 23.75
Date |  |  | RM1 |  | RM2 |  | RM3 |  | 5DOFx | 5DOFy | 5DOFz |  |  | AC Focus | Cam exp(ms) | Lidar Sig | Range(m) | .fits | .bmp | AC AZ | AC EL | X centroid | Y centroid |  | Range repeatability (m)
2023-01-25 |  | sMATF mirror | 21.68 | 21.66 | 21.53 | 25.27 | 24.15 | 23.75 | [9 empty] | 63026 |  | 0.0013 | -0.0034
2023-01-25 |  | sMATF mirror | 21.68 | 21.66 | 21.53 | 25.27 | 24.15 | 23.75 | [9 empty] | 63419 |  |  |  | 798.7 | 599.7
2023-01-25 |  | sMATF Retro | 21.68 | 21.66 | 21.53 | 25.27 | 24.15 | 23.75 | [7 empty] | 9 | 4.0322 | 63723 |  |  |  | 154 | 1044
2023-01-25 |  | PATB mirror, nominal rx and ry | 21.68 | 21.66 | 21.53 | 25.27 | 24.15 | 23.75 | 101.5 | 125.7 |  |  |  | 0,23 | [5 empty] | -0.0011 | -0.01
2023-01-25 |  | PATB primary LED, same as morning position | 21.68 | 21.66 | 21.53 | 25.27 | 24.15 | 23.75 | 88.69 | 108.4 |  |  |  | 1125 | [7 empty] | 797.7 | 600
2023-01-25 |  | PATB primary Retro, nomial position | 21.68 | 21.66 | 21.53 | 25.27 | 24.15 | 23.75 | 69.2 | 102.9 | [5 empty] | 17 | 3.4596 | 72009 |  |  |  | 219 | 932
Date |  |  | RM1 |  | RM2 |  | RM3 |  | 5DOFx | 5DOFy | 5DOFz |  |  | AC Focus | Cam exp(ms) | Lidar Sig | Range(m) | .fits | .bmp | AC AZ | AC EL | X centroid | Y centroid
2023-01-26 |  | sMATF mirror | 21.68 | 21.66 | 21.53 | 25.27 | 24.15 | 23.75 | [5 empty] | 0,25 |  |  |  | 54513 |  | 0.0024 | -0.0041
2023-01-26 |  | sMATF mirror | 21.68 | 21.66 | 21.53 | 25.27 | 24.15 | 23.75 | [5 empty] | 1,76 |  |  |  | 55109 |  |  |  | 797.6 | 600.4
2023-01-26 |  | sMATF Retro | 21.68 | 21.66 | 21.53 | 25.27 | 24.15 | 23.75 | [5 empty] | 1,77 |  | 8 | 4.0322 | 55504 |  |  |  | 150.5 | 1046.5
2023-01-26 |  | PATB mirror, nominal rx and ry | 21.68 | 21.66 | 21.53 | 25.27 | 24.15 | 23.75 | 101.503 | 125.705 |  |  |  | 0,25 |  |  |  | 61046 |  | -0.0001 | -0.0107 |  |  |  |  | Rx=0.1052, Ry=0.0608
2023-01-26 |  | PATB primary LED | 21.68 | 21.66 | 21.53 | 25.27 | 24.15 | 23.75 | 88.695 | 108.398 |  |  |  | 1131 |  |  |  | 62610 |  |  |  | 797.45 | 599.7
2023-01-26 |  | PATB primary Retro | 21.68 | 21.66 | 21.53 | 25.27 | 24.15 | 23.75 | 69.2 | 102.898 |  |  |  | 1138 |  | 23 | 3.4596 | 63001 |  |  |  | 220 | 933
2023-01-26 |  | sMATF mirror (5DOF stowed) | 21.68 | 21.66 | 21.53 | 25.27 | 24.15 | 23.75 | [5 empty] | 0,19 |  |  |  | 64738 |  | 0.0027 | -0.0065
2023-01-26 |  | sMATF mirror (5DOF stowed) | 21.68 | 21.66 | 21.53 | 25.27 | 24.15 | 23.75 | [5 empty] | 1,76 |  |  |  | 65016 |  |  |  | 795.6 | 600.7 |  |  | Slight shift, visible to the naked eye
2023-01-26 |  | sMATF Retro (5DOF stowed) | 21.68 | 21.66 | 21.53 | 25.27 | 24.15 | 23.75 | [5 empty] | 1,80 |  |  | 4.0322 | 65237 |  |  |  | 150 | 1048
Date |  |  | RM1 |  | RM2 |  | RM3 |  | 5DOFx | 5DOFy | 5DOFz |  |  | AC Focus | Cam exp(ms) | Lidar Sig | Range(m) | .fits | .bmp | AC AZ | AC EL | X centroid | Y centroid
2023-01-30 |  | sMATF mirror, H2 position | 21.68 | 21.66 | 21.53 | 25.27 | 24.15 | 23.75 | [5 empty] | 0,26 |  |  |  | 91359 |  | 0.0037 | -0.0058
2023-01-30 |  | sMATF mirror, pinhole | [11 empty] | 1,76 |  |  |  | 91634 |  |  |  | 795.7 | 601
2023-01-30 |  | sMATF Retro | [11 empty] | 1,81 |  | 10 | 4.0321 | 91852 |  |  |  | 150 | 1047 |  | -8.9e-05
2023-01-30 |  | PATB mirror, nominal rx and ry | 21.68 | 21.66 | 21.53 | 25.27 | 24.15 | 23.75 | 101.503 | 125.705 |  |  |  | 0,25 |  |  |  | 3943 |  | -0.0001 | -0.0104
2023-01-30 |  | PATB primary LED | 21.68 | 21.66 | 21.53 | 25.27 | 24.15 | 23.75 | 88.595 | 108.4 |  |  |  | 1131 |  |  |  | 4503 |  |  |  | 795.7 | 601.1
2023-01-30 |  | PATB primary Retro | 21.68 | 21.66 | 21.53 | 25.27 | 24.15 | 23.75 | 69.2 | 102.9 |  |  |  | 1138 |  | 17 | 3.4596 | 5015 |  |  |  | 220 | 931 |  | -3.2e-05
2023-01-30 |  | sMATF mirror (5DOF stowed) | 21.68 | 21.66 | 21.53 | 25.27 | 24.15 | 23.75 | [5 empty] | 0,19 |  |  |  | 10130 |  | 0.0024 | -0.0065
2023-01-30 |  | sMATF mirror (5DOF stowed) | 21.68 | 21.66 | 21.53 | 25.27 | 24.15 | 23.75 | [5 empty] | 1,76 |  |  |  | 10258 |  |  |  | 795.6 | 600.7
2023-01-30 |  | sMATF Retro (5DOF stowed) | 21.68 | 21.66 | 21.53 | 25.27 | 24.15 | 23.75 | [5 empty] | 1,80 |  | 7 | 4.0322 | 10448 |  |  |  | 148 | 1046 |  | 5.2e-05
Date |  |  | RM1 |  | RM2 |  | RM3 |  | 5DOFx | 5DOFy | 5DOFz |  |  | AC Focus | Cam exp(ms) | Lidar Sig | Range(m) | .fits | .bmp | AC AZ | AC EL | X centroid | Y centroid
2023-01-31 |  | sMATF mirror, H2 position | [11 empty] | 0,25 |  |  |  | 55125 |  | 0.0017 | -0.0048
2023-01-31 |  | sMATF mirror, pinhole | [11 empty] | 1,72 |  |  |  | 61422 |  |  |  | 796 | 599.9
2023-01-31 |  | sMATF Retro | [13 empty] | 9 | 4.0322 | 62238 |  |  |  | 148 | 1046 |  | 7.3e-05
2023-02-01 |  | PATB mirror, nominal rx=0.105 and ry=0.061 | [6 empty] | 101.5 | 125.7 |  |  |  | 0,25 |  |  |  | 91609 |  | 0.0006 | 0.012
2023-02-01 |  | PATB primary LED | [6 empty] | 88.7 | 108.5 |  |  |  | 1133 |  |  |  | 92814 |  |  |  | 795.8 | 599.8
2023-02-01 |  | PATB primary Retro | [6 empty] | 69.2 | 102.9 |  |  |  | 1135 |  |  | 3.4595 | 93446 |  |  |  | 218 | 934 |  | -5.6e-05
Date |  |  | RM1 |  | RM2 |  | RM3 |  | 5DOFx (mm) | 5DOFy (mm) | 5DOFz (mm) | 5DOF Rx (deg) | 5DOF Ry (deg) | AC Focus | Cam exp(ms) | Lidar Sig | Range (m) | .fits | .bmp | AC AZ | AC EL | X centroid | Y centroid
2023-02-02 |  | sMATF mirror, H2 position | [6 empty] | 250 | -65 | -60.0012 | -0.0003 | 0.0003 | 0,25 |  |  |  | 41848 |  | 0.0048 | -0.0048
2023-02-02 |  | sMATF mirror, pinhole, H2 position | [6 empty] | 250 | -65 | -60.0012 | -0.0003 | 0.0003 | 1,73 |  |  |  | 43107 |  |  |  | 795.724 | 601.921
2023-02-02 |  | sMATF Retro, H2 position | [6 empty] | 250 | -65 | -60.0012 | -0.0003 | 0.0003 | 1,79 |  |  | 4.0321 | 43424 |  |  |  | 151 | 1049 |  | -79 | 1.1655
2023-02-02 |  | PATB mirror, nominal rx=0.105 and ry=0.061 | [6 empty] | 101.501 | 125.703 | 0.0034 | 0.10547 | 0.06164 | 0,24 |  |  |  | 50814 |  | -0.0001 | -0.011
2023-02-02 |  | PATB primary LED | [6 empty] | 88.596 | 108.399 | 0.0005 | 0.10556 | 0.0617 | 1135 |  |  |  | 52350 |  |  |  | 795.707 | 601.771
2023-02-02 |  | PATB primary Retro | [6 empty] | 69.2002 | 102.898 | -0.00098 | 0.10596 | 0.06107 | 1139 |  |  | 3.4596 | 53120 |  |  |  | 220 | 931 |  | 2.8e-05
Note: for this check we are purposely NOT optimizing RMs during sMATF measurement or 5DOF positions during PATB measurement Turns out, we wouldn't have needed to do any adjustments anyway because things were in tolerance with nominal positions | [25 empty] | For ADM on Feb3:
2023-02-03 |  | sMATF mirror, H2 position | [6 empty] | -2147.5 | -65 | -60 | 0 | 0 | 0,20 |  |  |  | 33856 |  | 0.002 | -0.0076 |  |  |  |  | Keep RMs the same, go to H2 position
2023-02-03 |  | sMATF mirror, pinhole, H2 position | [11 empty] | 1,79 |  |  |  | 34915 |  |  |  | 795.8 | 599.7 |  |  | Measure Az and El and centroid (do not center or minimize Az, el)
2023-02-03 |  | sMATF Retro, H2 position | [11 empty] | 1,88 |  |  | 4.0322 | 35549 |  |  |  | 151 | 1046 |  | 123 | Range if possible
2023-02-03 |  | PATB mirror, nominal rx=0.105 and ry=0.061 | [6 empty] | 101.502 | 125.703 | -0.003 | 0.1055 | 0.06149 | 0,30 |  |  |  | 42244 |  | -0.0025 | -0.0104 |  |  |  |  | Go to PATB with nominal/previous 5DOF positions measure the Az, El, centroid position and range if possible.
2023-02-03 |  | PATB primary LED | [6 empty] | 88.501 | 108.397 | -0.002 | 0.1056 | 0.0615 | 1137 |  |  |  | 43354 |  |  |  | 794.5 | 601.6 |  |  | (start nominal ADM ops here) Go back to H2 and use RMs to bring Az, El and centroid into alignment and move rm3 if needed to range.
2023-02-03 |  | PATB primary Retro | [6 empty] | 69.202 | 102.897 | -0.003 | 0.1056 | 0.0616 | 1141 |  |  | 3.4596 | 44338 |  |  |  | 219 | 931 |  |  | Go to PATB and adjust 5DOF to match sMATF az, el and centroid as needed. Adjust 5DOF to range as needed.
2023-02-03 |  | PATA mirror | [6 empty] | -76.103 | 125.701 | 0.002 | 0.0928 | 0.061 | 0,23 |  |  |  | 51403 |  | 0.0013 | 0.0001 |  |  |  |  | Go to PATA and repeat 5DOF adjustment to match sMATF az, el and centroid as needed. Adjust 5DOF to range as needed.
2023-02-03 |  | PATA primary LED | [6 empty] | -60.395 | 108.399 | 0.0001 | 0.0929 | 0.0611 | 1132 |  |  |  | 55628 |  |  |  | 796.6 | 599.6
2023-02-03 |  | PATA primary Retro | [6 empty] | -103.996 | 102.701 | -0.001 | 0.0929 | 0.0612 | 1132 |  |  | 3.4596 | 60607 |  |  |  | 215 | 935
2023-02-03 |  | PATA retro & LED in the same image | 21.68 | 21.66 | 21.53 | 25.27 | 24.15 | 23.75 | [9 empty] | can't see both the LED and retro at the same time? At least not when we are at the retro target
Date |  |  | RM1 |  | RM2 |  | RM3 |  | 5DOFx (mm) | 5DOFy (mm) | 5DOFz (mm) | 5DOF Rx (deg) | 5DOF Ry (deg) | AC Focus | Cam exp(ms) | Lidar Sig | Range (m) | .fits | .bmp | AC AZ | AC EL | X centroid | Y centroid
2023-02-05 |  | sMATF mirror, H2 position | 21.66 | 23.63 | 21.44 | 24.89 | 24.15 | 23.75 | -2147.48 | -65.001 | -60.0047 | -0.0008 | 0.00031 | 0,23 |  |  |  | 80159 |  | 0.0037 | -0.0065
2023-02-05 |  | sMATF mirror, pinhole, H2 position | 21.66 | 23.63 | 21.44 | 24.89 | 24.15 | 23.75 | -2147.48 | -65.001 | -60.0047 | -0.0008 | 0.00031 | 1,77 |  |  |  | 75615 |  |  |  | 792.6 | 604.4
2023-02-05 |  | sMATF Retro, H2 position | 21.66 | 23.63 | 21.44 | 24.89 | 24.15 | 23.75 | -2147.48 | -65.001 | -60.0047 | -0.0008 | 0.00031 | 1,84 |  |  | 4.0322 | 80641 |  |  |  | 149 | 1048
2023-02-05 |  | PATB mirror | 21.66 | 23.63 | 21.44 | 24.89 | 24.15 | 23.75 | 101.5 | 125.7 | 0.0018 | 0.1055 | 0.0612 | 0,24 |  |  |  | 82833 |  | -0.0022 | -0.0107
2023-02-05 |  | PATB primary LED | 21.66 | 23.63 | 21.44 | 24.89 | 24.15 | 23.75 | 88.502 | 108.35 | 0.0005 | 0.1056 | 0.0612 | 1139 |  |  |  | 84611 |  |  |  | 792.7 | 603.7
2023-02-05 |  | PATB primary Retro | 21.66 | 23.63 | 21.44 | 24.89 | 24.15 | 23.75 | 69.2 | 102.898 | 0.0005 | 0.1057 | 0.0612 | 1149 |  |  | 3.4596 | 85435 |  |  |  | 219 | 933
2023-02-05 |  | PATB primary Retro @ sMATF pinhole | 21.66 | 23.63 | 21.44 | 24.89 | 24.15 | 23.75 | 100.5 | 100.9 | 0.0005 | 0.1057 | 0.0612 | 1145 |  |  |  | 3719 |  |  |  | 792 | 612 |  |  | VERY uncertain. +/-500 microns?
2023-02-05 |  | PATB redundant Retro | 21.66 | 23.63 | 21.44 | 24.89 | 24.15 | 23.75 | 92.911 | 144.9 | 0.0012 | 0.1054 | 0.0613 | 1134 |  |  | 3.4676 | 4541 |  |  |  | 217 | 932
2023-02-05 |  | PATB redundant Retro @ sMATF pinhole | 21.66 | 23.63 | 21.44 | 24.89 | 24.15 | 23.75 | 125.5 | 143.2 | 0.0005 | 0.106 | 0.061 | 1142 |  |  |  | 11744 |  |  |  | 806 | 624 |  |  | VERY uncertain. +/-500 microns?
2023-02-05 |  | PATA mirror | 21.66 | 23.63 | 21.44 | 24.89 | 24.15 | 23.75 | -76.1 | 125.7 | 0 | 0.09317 | 0.06063 | 0,19 |  |  |  | 12504 |  | 0.0006 | 0.0014
2023-02-05 |  | PATA primary LED | 21.66 | 23.63 | 21.44 | 24.89 | 24.15 | 23.75 | -60.6012 | 108.3 | 0.0008 | 0.0926 | 0.06063 | 1136 |  |  |  | 14219 |  |  |  | 792 | 603.9
2023-02-05 |  | PATA primary Retro | 21.66 | 23.63 | 21.44 | 24.89 | 24.15 | 23.75 | -103.995 | 102.7 | 0.0005 | 0.0927 | 0.0613 | 1141 |  |  | 3.4596 | 15426 |  |  |  | 215 | 933
2023-02-05 |  | PATA primary Retro @ sMATF pinhole | 21.66 | 23.63 | 21.44 | 24.89 | 24.15 | 23.75 | [17 empty] | Not visible; intensity drops to zero while translating to center.
2023-02-05 |  | PATA redundant Retro | 21.66 | 23.63 | 21.44 | 24.89 | 24.15 | 23.75 | [17 empty] | Not visible
2023-02-05 |  | PATA redundant Retro @ sMATF pinhole | 21.66 | 23.63 | 21.44 | 24.89 | 24.15 | 23.75
Date |  |  | RM1 |  | RM2 |  | RM3 |  | 5DOFx (mm) | 5DOFy (mm) | 5DOFz (mm) | 5DOF Rx (deg) | 5DOF Ry (deg) | AC Focus | Cam exp(ms) | Lidar Sig | Range (m) | .fits | .bmp | AC AZ | AC EL | X centroid | Y centroid
2023-02-06 |  | PATA primary Retro | 21.66 | 23.63 | 21.44 | 24.89 | 24.15 | 23.75 | -104 | 102.7 |  | 0.093 | 0.061 | 1140 |  |  |  | 13108 |  |  |  | 214 | 935 |  |  | 51.3128 | 20.5251 | Shift PAT retro image to center
2023-02-06 |  |  | 21.66 | 23.63 | 21.44 | 24.89 | 24.15 | 23.75 | -106.5 | 102.7 |  | 0.093 | 0.061 |  |  |  |  | 14220 |  |  |  | 171 | 963 |  | 30.6454 | -17.6857 | Scale in PAT plane C(pix/mm) | RM3a | RM3b | x | y
2023-02-06 |  | Steer Rm3 to illuminate center FOV with Lidar, and translate 5Dof x-y to move PATA primary Retro to the center | 21.66 | 23.63 | 21.44 | 24.89 | 24.4 | 23.45 | -71.86 | 100 |  | 0.093 | 0.061 | 1140 |  |  |  | 41128 |  |  |  | 800 | 600 |  | -68.6 | 674.997 | 20.9281 | 0.25 | -0.3 | 32.14 | -2.7
2023-02-06 |  | Better focus Az (will cause small unwanted image shift) | 21.66 | 23.63 | 21.44 | 24.89 | 24.4 | 23.45 | -71.86 | 100 |  | 0.093 | 0.061 | 1143 |  |  |  | 41549 |  |  |  | 802 | 598 |  | 101.1
2023-02-06 |  | PATB primary Retro @ AC image center | 21.66 | 23.63 | 21.44 | 24.89 | 24.4 | 23.45 | 101 | 100.2 |  | 0.106 | 0.06 | 1143 |  |  |  | 44536 |  |  |  | 800 | 600
2023-02-06 |  | PATB primary Retro @ AC image corner by restoring RM3 | 21.66 | 23.63 | 21.44 | 24.89 | 24.15 | 23.75 | 69.2 | 102.9 |  | 0.106 | 0.06 | 1143 |  |  |  | 45223 |  |  |  | 214 | 932
2023-02-06 |  | SMATF Retro | 21.66 | 23.63 | 21.44 | 24.89 | 24.12 | 23.77 | -2150 | -65 |  | 0 | 0 | 1,79 |  | 43 | 4.0323 | 55910 |  |  |  | 146 | 1046 |  | 5.8e-05
2023-02-06 |  | Chips Retro | 21.66 | 23.63 | 21.44 | 24.89 | 24.62 | 23.25 | -2150 | -65 |  | 0 | 0 | 1,52 |  | 21 | 4.7631 | 63201 |  |  |  | 1278 | 220
Date |  |  | RM1 |  | RM2 |  | RM3 |  | 5DOFx (mm) | 5DOFy (mm) | 5DOFz (mm) | 5DOF Rx (deg) | 5DOF Ry (deg) | AC Focus | Cam exp(ms) | Lidar Sig | Range (m) | .fits | .bmp | AC AZ | AC EL | X centroid | Y centroid
2023-02-07 |  | Chips Retro | 21.66 | 23.63 | 21.44 | 24.89 | 24.62 | 23.25 | [5 empty] | 1,52 |  |  | 4.7632 | 11639 |  |  |  | 1285 | 214
2023-02-07 |  | SMATF Retro | 21.66 | 23.63 | 21.44 | 24.89 | 24.12 | 23.77 | [5 empty] | 1,79 |  |  | 4.0323 | 12030 |  |  |  | 145.3 | 1045.5 |  | 1.8e-05
2023-02-07 |  | sMATF Mirror | 21.71 | 23.63 | 21.44 | 25.29 | 24.12 | 23.77 | [11 empty] | -0.02 | -0.024 | 800 | 599
2023-02-07 |  | sMATF Mirror, Rm homing and optimization | 21.71 | 23.42 | 21.13 | 25.3 | 24.12 | 23.77 | [5 empty] | 0,25 |  |  |  | 25132 |  | 0.001 | 0.003
2023-02-07 |  |  | 21.71 | 23.42 | 21.13 | 25.3 | 24.12 | 23.77 | [5 empty] | 1,79 |  |  |  | 25413 |  |  |  | 803 | 616
2023-02-07 |  | SMATF Retro | 21.71 | 23.42 | 21.13 | 25.3 | 24.12 | 23.77 | 250 | -65 | -60 | 0 | 0 |  |  | 25 | 4.0323 | 30230 |  |  |  | 146.4 | 1059.5
2023-02-07 |  | PATB Mirror | 21.71 | 23.42 | 21.13 | 25.3 | 24.12 | 23.77 | 101.502 | 125.706 | 0.003 | 0.0875 | 0.0598 | 0,25 |  |  |  | 33909 |  | 0.0065 | 0.0067
2023-02-07 |  | PATB primary LED | 21.71 | 23.42 | 21.13 | 25.3 | 24.12 | 23.77 | 88.8445 | 108.249 | -0.001 | 0.0875 | 0.0598 | 1128 |  |  |  | 35137 |  |  |  | 802 | 616
2023-02-07 |  | PATB Pri Retro | 21.71 | 23.42 | 21.13 | 25.3 | 24.12 | 23.77 | 67.6975 | 103.398 | -0.001 | 0.0875 | 0.0598 | 1140 |  | 18 | 3.4597 | 40023 |  |  |  | 185 | 950 |  | 0.000133
2023-02-07 |  | PATB Pri Retro at Center, optimal focus | 21.71 | 23.42 | 21.13 | 25.3 | 24.4 | 23.45 | 100.9 | 100.4 | -0.002 | 0.0875 | 0.0598 | 1145 |  |  |  | 42600 |  |  |  | 798 | 614
2023-02-07 |  | PATB Pri Retro | 21.71 | 23.42 | 21.13 | 25.3 | 24.12 | 23.77 | 67.6975 | 103.397 | -0.001 | 0.0875 | 0.0598 |  |  |  | 3.4597 | [7 empty] | 1e-06
2023-02-07 |  | PATA mirror | 21.71 | 23.42 | 21.13 | 25.3 | 24.12 | 23.77 | -76.101 | 125.701 | -0.001 | 0.07 | 0.05 | 0,25 |  |  |  | 45119 |  | 0.013 | 0.008
2023-02-07 |  | PATA primary LED | 21.71 | 23.42 | 21.13 | 25.3 | 24.12 | 23.77 | -60.282 | 108.199 | -0.002 | 0.07 | 0.05 | 1137 |  |  |  | 50425 |  |  |  | 803 | 617 |  | 0.57269
2023-02-07 |  | PATA primary Retro | 21.71 | 23.42 | 21.13 | 25.3 | 24.12 | 23.77 | -105.494 | 102.701 | 0 | 0.7032 | 0.04991 | 1137 |  |  | 3.4596 | 51518 |  |  |  | 189 | 962
Date |  |  | RM1 |  | RM2 |  | RM3 |  | 5DOFx (mm) | 5DOFy (mm) | 5DOFz (mm) | 5DOF Rx (deg) | 5DOF Ry (deg) | AC Focus | Cam exp(ms) | Lidar Sig | Range (m) | .fits | .bmp | AC AZ | AC EL | X centroid | Y centroid
2023-02-08 |  | sMATF mirror, reoptimized after SALSA mirror adjustment | 20.7032 | 23.43 | 21.15 | 24.07 | 24.12 | 23.77 | [5 empty] | 0,18 |  |  |  | 4628 |  | 0.0008 | 0.0004
2023-02-08 |  | sMATF mirror, pinhole, H2 position | 20.7032 | 23.43 | 21.15 | 24.07 | 24.12 | 23.77 | [5 empty] | 1,79 |  |  |  | 4937 |  |  |  | 793.1 | 602.1
2023-02-08 |  | sMATF Retro, H2 position | 20.7032 | 23.43 | 21.15 | 24.07 | 24.12 | 23.77 | [5 empty] | 1,80 |  |  | 4.0316 | 5907 |  |  |  | 148 | 1051 |  | -0.0007
2023-02-08 |  | PATB mirror | 20.7032 | 23.43 | 21.15 | 24.07 | 24.12 | 23.77 | 101.501 | 125.705 | 0.0006 | 0.08677 | 0.0597 | 0,19 |  |  |  | 11519 |  | 0.0062 | 0.0039
2023-02-08 |  | PATB primary LED | 20.7032 | 23.43 | 21.15 | 24.07 | 24.12 | 23.77 | 88.8307 | 108.199 | 0.0004 | 0.0875 | 0.05975 | 1131 |  |  |  | 12421 |  |  |  | 792.7 | 601.7
2023-02-08 |  | PATB primary Retro | 20.7032 | 23.43 | 21.15 | 24.07 | 24.12 | 23.77 | 67.6959 | 103.397 | 0.0004 | 0.08753 | 0.05974 | 1140 |  |  | 3.4591 | 13009 |  |  |  | 179 | 941
2023-02-08 |  | PATB primary Retro @ sMATF pinhole | 20.7032 | 23.43 | 21.15 | 24.07 | 24.4 | 23.45 | 100.905 | 100.4 | 0.0008 | 0.08755 | 0.05976 | 1140 |  |  |  | 13906 |  |  |  | 793 | 603
2023-02-08 |  | PATB redundant LED | 20.7032 | 23.43 | 21.15 | 24.07 | 24.12 | 23.77 | 111.9 | 148.601 | 0.0004 | 0.08719 | 0.05976 | 1131 |  |  |  | 20329 |  |  |  | 792.6 | 601
Date |  |  | RM1 |  | RM2 |  | RM3 |  | 5DOFx (mm) | 5DOFy (mm) | 5DOFz (mm) | 5DOF Rx (deg) | 5DOF Ry (deg) | AC Focus | Cam exp(ms) | Lidar Sig | Range (m) | .fits | .bmp | AC AZ | AC EL | X centroid | Y centroid
2023-02-10 |  | 5DOF @ H2, prior to range repeatability of sMATF retro | 20.7 | 23.43 | 21.15 | 24.07 | 24.12 | 23.77 | -2147.48 | -65.001 | -60.0054 | -0.0009 | 0.0004 | 1,80 |  | 25 | many measurements performed here. See range files
2023-02-10 |  | 5DOF @ H2, prior to range repeatability of Chips retro | 20.7 | 23.43 | 21.15 | 24.07 | 24.64 | 23.24 | -2147.48 | -65.001 | -60.0054 | -0.0009 | 0.0004 | 1,52 |  | 7 | many measurements performed here. See range files
2023-02-10 |  | PATB primary Retro (at nominal position) | 20.703 | 23.43 | 21.15 | 24.07 | 24.12 | 23.77 | 67.69 | 103.399 | -0.004 | 0.0875 | 0.0599 |  |  |  | many measurements performed here. See range files
2023-02-10 |  | PATB primary LED (0,0) | 20.703 | 23.43 | 21.15 | 24.07 | 24.12 | 23.77 | 88.8293 | 108.199 | 0 | 0.0875 | 0.0599 | 1131 |  |  |  | 14626 |  |  |  | 829.7 | 601.9
2023-02-10 |  | PATB primary LED (9,0) | 20.703 | 23.43 | 21.15 | 24.07 | 24.12 | 23.77 | 97.8297 | 108.199 | 0 | 0.0875 | 0.0599 | 1131 |  |  |  | 14808 |  |  |  | 985.1 | 494.8 |  |  |  |  | Scale (pix/mm) |  |  |  | Clocking (deg)
2023-02-10 |  | PATB primary LED (-9,0) | 20.703 | 23.43 | 21.15 | 24.07 | 24.12 | 23.77 | 79.843 | 108.199 | 0 | 0.0875 | 0.0599 | 1131 |  |  |  | 15052 |  |  |  | 674.6 | 708 |  |  |  | 376.649 | #DIV/0! |  | -310.5 | 213.2 | 34.4748
2023-02-10 |  | PATB primary LED (0,9) | 20.703 | 23.43 | 21.15 | 24.07 | 24.12 | 23.77 | 88.83 | 117.2 | -0.0005 | 0.0875 | 0.0599 | 1131 |  |  |  | 15315 |  |  |  | 724.9 | 449.7
2023-02-10 |  | PATB primary LED (0,-9) | 20.703 | 23.43 | 21.15 | 24.07 | 24.12 | 23.77 | 88.8302 | 99.2 | 0.0001 | 0.0875 | 0.0599 | 1131 |  |  |  | 15504 |  |  |  | 934 | 754.7
Date |  |  | RM1 |  | RM2 |  | RM3 |  | 5DOFx (mm) | 5DOFy (mm) | 5DOFz (mm) | 5DOF Rx (deg) | 5DOF Ry (deg) | AC Focus | Cam exp(ms) | Lidar Sig | Range (m) | .fits | .bmp | AC AZ | AC EL | X centroid | Y centroid
2023-02-14 |  | Chips Retro(RMs roughed) | 20.702 | 23.427 | 21.138 | 24.063 | 24.63 | 23.23 | -2147.48 | -65.001 | -60.0054 | -0.0009 | 0.0004 |  |  | 7 | 4.7624 | [8 empty] | 13:37 | 305.1 | 301.2 | 301 | 302.7 | 306.2 | 300.6 | 302.7 | disabled
2023-02-14 |  | sMATF Retro(RMs roughed) | 20.702 | 23.427 | 21.138 | 24.063 | 24.12 | 23.77 | -2147.48 | -65.001 | -60.0054 | -0.0009 | 0.0004 |  |  | 36 | 4.0316 | [8 empty] | 13:40 | 305.4 | 301.3 | 301.1 | 303 | 307.3 | 301 | 303.3 | disabled
2023-02-14 |  | sMATF Retro (starting point) | 20.7 | 23.43 | 21.15 | 24.07 | 24.12 | 23.77 | [11 empty] | 0.0228 | 0.0285 | 800.77 | 606.69
2023-02-14 |  | sMATF Retro (after re-homing RM1) | 20.7 | 23.43 | 21.15 | 24.07 | 24.12 | 23.77 | [11 empty] | 0.0402 | 0.0133 | 789 | 619.7
2023-02-14 |  | sMATF Retro (after re-homing RM1 again) | 20.7 | 23.43 | 21.15 | 24.07 | 24.12 | 23.77 | [11 empty] | 0.0385 | 0.0112 | 785.8 | 620.7 |  | Plate scale: 1 pixel = 60 microns.
2023-02-14 |  | sMATF Retro (after re-homing RM2) | 20.7 | 23.43 | 21.15 | 24.07 | 24.12 | 23.77 | [11 empty] | 0.0284 | 0.0167 | 789.8 | 612.7
2023-02-14 |  | sMATF Retro (after jogging RM2 positive then negative .02mm) | 20.7 | 23.43 | 21.15 | 24.07 | 24.12 | 23.77 | [11 empty] | 0.034 | 0.0133 | 788.6 | 615.7
2023-02-14 |  | sMATF Retro (after jogging RM2 negative then positive .02mm) | 20.7 | 23.43 | 21.15 | 24.07 | 24.12 | 23.77 | [11 empty] | 0.0239 | 0.0209 | 794.7 | 607.7 |  |  | 15:00 | 305.5 | 301.4 | 301.2 | 303 | 307.5 | 300.9 | 303.5 | disabled
2023-02-14 |  | sMATF Retro (after jogging RM2 positive then negative .05mm) | 20.7 | 23.43 | 21.15 | 24.07 | 24.12 | 23.77 | [11 empty] | 0.035 | 0.0133 | 787.9 | 616.6
2023-02-14 |  | sMATF Retro (after jogging RM2 negative then positive .05mm) | 20.7 | 23.43 | 21.15 | 24.07 | 24.12 | 23.77 | [11 empty] | 0.0242 | 0.0209 | 794.7 | 607.8
2023-02-14 |  | sMATF Retro (after re-homing RM2) | 20.7 | 23.43 | 21.15 | 24.07 | 24.12 | 23.77 | [11 empty] | 0.0239 | 0.016 | 791 | 606.7
2023-02-14 |  | *Changed the ADM RM backlash correction parameter from 0.02 mm to 0.1 mm
2023-02-14 |  | sMATF retro (after re-homing RM1 and RM2) | 20.7 | 23.43 | 21.15 | 24.07 | 24.12 | 23.77 | [11 empty] | 0.026 | 0.0081 | 784.7 | 607.8
2023-02-14 |  | sMATF Retro (after jogging RM1 positive then negative .1mm) | 20.7 | 23.43 | 21.15 | 24.07 | 24.12 | 23.77 | [11 empty] | 0.0162 | 0.0258 | 796.8 | 602.8
2023-02-14 |  | sMATF Retro (after jogging RM1 negative then positive .1mm) | 20.7 | 23.43 | 21.15 | 24.07 | 24.12 | 23.77 | [11 empty] | 0.0284 | 0.0081 | 784.6 | 610.8
2023-02-14 |  | sMATF Retro (after jogging RM2 negative then positive .1mm) | [17 empty] | 0.0294 | 0.007 | 783.7 | 611.8 |  |  | 15:33 | 305.6 | 301.4 | 301.2 | 303 | 307.6 | 301 | 303.6 | disabled
2023-02-14 |  | sMATF Retro (after homing RM2) | [19 empty] | 787.8 | 611.6
2023-02-14 |  | sMATF Retro (after homing RM2 again) | [19 empty] | 788.7 | 609.9
2023-02-14 |  | sMATF Retro (after homing RM2 again) | [19 empty] | 788.6 | 610.8
2023-02-14 |  | sMATF Retro (after jogging RM2 negative then positive .1mm) | [19 empty] | 789.7 | 609.8
2023-02-14 |  | sMATF Retro (after jogging RM2 negative then positive .1mm) | [19 empty] | 790.4 | 609.66
2023-02-14 |  | sMATF Retro (after jogging RM2 negative then positive .1mm) | [19 empty] | 790.5 | 609.6
2023-02-14 |  | Found a way to null ADM with RM1,2 | 20.57 | 23.83 | 21.701 | 23.94 | 24.12 | 23.77 | [9 empty] | 81032 |  |  |  | 807.7 | 601.6
2023-02-14 |  |  | 20.57 | 23.83 | 21.701 | 23.94 | 24.12 | 23.77 | [8 empty] | 4.0316 | 81341 |  | 0.005 | 0.004
2023-02-14 |  | PATB Pri Retro nominal position | [9 empty] | 0.087 | 0.0595 |  |  |  | 3.4591
2023-02-14 |  | PATB mirror | [6 empty] | 101.5 | 125.7 | 0 | 0.09 | 0.06 | 0,23 |  |  |  | 11249 |  | -0.005 | -0.008
2023-02-14 |  | PATB Pri Retro, new RX Ry | [6 empty] | 67.7 | 103.4 | 0 | 0.09 | 0.06 |  |  | 17 | 3.459
2023-02-14 |  | PATB LED pri | [6 empty] | 88.49 | 108.4 | 0 | 0.09 | 0.06 | 1131 |  |  |  | 13156 |  |  |  | 808 | 602
Date |  |  | RM1 |  | RM2 |  | RM3 |  | 5DOFx (mm) | 5DOFy (mm) | 5DOFz (mm) | 5DOF Rx (deg) | 5DOF Ry (deg) | AC Focus | Cam exp(ms) | Lidar Sig | Range (m) | .fits | .bmp | AC AZ | AC EL | X centroid | Y centroid
2023-02-15 |  | PATB mirror (at start of day) | [6 empty] | 101.5 | 125.707 | 0 | 0.09 | 0.06 | 0,20 | [5 empty] | -0.0042 | -0.0107 |  |  |  |  | 10:03 | 302.2 | 300.4 | 300.1 | 300.9 | 302.4 | 299.5 | 300.1 | disabled
2023-02-15 |  | PATB mirror (after correction 1) | [6 empty] | 101.5 | 125.707 | 0 | 0.1068 | 0.0665 | 0,20 |  |  |  | 3043 |  | 0.0048 | 0.0039
2023-02-15 |  | PATB primary LED (0,0) | [6 empty] | 88.5837 | 108.298 | -0.0015 | 0.1069 | 0.0666 | 1124 |  |  |  | 4216 |  |  |  | 807.7 | 602.6
2023-02-15 |  | PATB primary LED (9,0) | [6 empty] | 97.5775 | 108.298 | -0.0014 | 0.10689 | 0.0666 | 1124 |  |  |  | 4453 |  |  |  | 961.7 | 492.8 |  | 154 | -109.8
2023-02-15 |  | PATB primary LED (-9,0) | [6 empty] | 79.592 | 108.298 | -0.0018 | 0.10689 | 0.0666 | 1124 |  |  |  | 4614 |  |  |  | 653.7 | 710.9 |  |  | 189.135
2023-02-15 |  | PATB primary LED (0,9) | [6 empty] | 88.5782 | 117.297 | -0.0015 | 0.10686 | 0.0666 | 1124 |  |  |  | 4744 |  |  |  | 700.7 | 450.9 |  |  | 0.047585
2023-02-15 |  | PATB primary LED (0,-9) | [6 empty] | 88.5784 | 99.2977 | -0.0016 | 0.10694 | 0.0666 | 1124 |  |  |  | 4916 |  |  |  | 914.6 | 752.9
2023-02-15 |  | PATB primary LED (0,0) (again) | [6 empty] | 88.5784 | 108.297 | -0.00165 | 0.10692 | 0.0666 | 1124 |  |  |  | 5051 |  |  |  | 807.6 | 601.9 |  |  | 10:51 | 303.6 | 300.7 | 300.5 | 301.7 | 305.7 | 299.9 | 301.7 | disabled
2023-02-15 |  | PATB primary LED (repeats) | [6 empty] | 88.5785 | 108.297 | -0.0017 | 0.1069 | 0.0666 | 1124 |  |  |  | 5532 |  |  |  | 807.7 | 601.9 |  |  | 10:56
2023-02-15 |  | PATB primary LED (repeats) | [6 empty] | 88.5786 | 108.297 | -0.0017 | 0.1069 | 0.0666 | 1124 |  |  |  | 5957 |  |  |  | 807.6 | 602.5 |  |  | 11:00
2023-02-15 |  | PATB primary LED (repeats) | [6 empty] | 88.5787 | 108.297 | -0.0016 | 0.1069 | 0.0666 | 1124 |  |  |  | 10723 |  |  |  | 807.6 | 601.8 |  |  | 11:08
2023-02-15 |  | PATB primary LED (repeats) | [6 empty] | 88.5788 | 108.297 | -0.0015 | 0.1069 | 0.0666 | 1124 |  |  |  | 11347 |  |  |  | 807.6 | 601.8 |  |  | 11:13
2023-02-15 |  | PATB primary LED (repeats) | [6 empty] | 88.5788 | 108.297 | -0.0015 | 0.1069 | 0.0666 | 1124 |  |  |  | 11638 |  |  |  | 807.6 | 601.8 |  |  | 11:17
2023-02-15 |  | PATB primary LED (repeats) | [6 empty] | 88.5788 | 108.297 | -0.0016 | 0.1069 | 0.0666 | 1124 |  |  |  | 12504 |  |  |  | 807.7 | 601.7 |  |  | 11:26 | 304.3 | 300.8 | 300.7 | 302.1 | 306.3 | 300.2 | 302.3 | disabled
2023-02-15 |  | PATB primary LED (repeats) | [6 empty] | 88.579 | 108.297 | -0.0017 | 0.1069 | 0.0666 | 1124 |  |  |  | 13040 |  |  |  | 807.7 | 601.7 |  |  | 11:31
2023-02-15 |  | PATB primary LED (repeats) | [6 empty] | 88.579 | 108.297 | -0.0016 | 0.1069 | 0.0666 | 1124 |  |  |  | 13551 |  |  |  | 807.7 | 601.7 |  |  | 11:36
2023-02-15 |  | PATB primary LED (repeats) | [6 empty] | 88.579 | 108.297 | -0.0017 | 0.1069 | 0.0666 | 1124 |  |  |  | 14334 |  |  |  | 807.7 | 601.6 |  |  | 11:44 | 304.5 | 301 | 300.8 | 302.3 | 306.5 | 300.4 | 302.5 | disabled
2023-02-15 |  | PATB mirror (re-check) | [6 empty] | 101.499 | 125.708 | 0.0002 | 0.10678 | 0.0666 | 0,30 |  |  |  | 22418 |  | 0.002 | 0.0046
2023-02-15 |  | sMATF mirror (Re-check, RMs unchanged from before) | 20.57 | 23.83 | 21.701 | 23.94 | 24.12 | 23.77 | -2147.48(+250) | -65.001 | -60.0067 | -0.0009 | -0.0003 | 0,20 |  |  |  | 23832 |  | 0.0037 | 0.0015 |  |  |  |  | 12:30 | 304.9 | 301.1 | 300.9 | 302.6 | 307 | 300.6 | 302.9 | disabled
2023-02-15 |  | PATB primary Retro | [6 empty] | 67.698 | 103.398 | -0.0002 | 0.1066 | 0.0667 |  |  | 16 | 3.4591 | 30211 | [6 empty] | laser on at ~12:50 | 13:01
2023-02-15 |  | PATB primary Retro (centered at pinhole) | [6 empty] | 100.71 | 100.401 | 0.0003 | 0.10658 | 0.06675 | 1140 |  |  |  | 30947 |  |  |  | 810 | 603 |  | powered off camera, LEDs, focus, leica, RM controller to help ADM cool a bit | 13:17 | 305.2 | 301.2 | 301 | 302.8 | 305 | 300.7 | 302.9 | disabled
|  | metrology happening here | [23 empty] | 13:54 | 304.4 | 301.1 | 300.9 | 302.1 | 302.2 | 300.4 | 302.4 | disabled
[26 empty] | 14:25 | 303.7 | 300.9 | 300.8 | 301.6 | 301.5 | 300.1 | 301.8 | disabled
Date |  |  | RM1 |  | RM2 |  | RM3 |  | 5DOFx (mm) | 5DOFy (mm) | 5DOFz (mm) | 5DOF Rx (deg) | 5DOF Ry (deg) | AC Focus | Cam exp(ms) | Lidar Sig | Range (m) | .fits | .bmp | AC AZ | AC EL | X centroid | Y centroid
2023-02-16 |  | CHIPs retro | 20.703 | 23.43 | 21.15 | 24.07 | 24.66 | 23.23 | -2147.48 | -65.001 | -60.0057 | -0.000926 | -0.0001995 |  |  | 7 | 4.7623 | [8 empty] | 10:03 | 302.2 | 300.4 | 300.1 | 300.9 | 302.4 | 299.5 | 300.1 | disabled
2023-02-16 |  | sMATF retro | 20.703 | 23.43 | 21.15 | 24.07 | 24.13 | 23.77 | -2147.48 | -65.001 | -60.0057 | -0.000926 | -0.0001995 |  |  | 26 | 4.0315
2023-02-16 |  | sMATF mirror (after RM optimization) | 20.521 | 23.806 | 21.691 | 23.884 | 24.13 | 23.77 | -2147.48 | -65.001 | -60.0057 | -0.000926 | -0.0001995 | 0,20 |  |  |  | 22626 |  | 0.0013 | 0.0028
2023-02-16 |  | sMATF LED (after RM optimization) | 20.521 | 23.806 | 21.691 | 23.884 | 24.13 | 23.77 | -2147.48 | -65.001 | -60.0057 | -0.000926 | -0.0001995 | 1,80 |  |  |  | 22759 |  |  |  | 807.8 | 601.5
2023-02-16 |  | sMATF retro (after RM optimization) | 20.521 | 23.806 | 21.691 | 23.884 | 24.13 | 23.77 | -2147.48 | -65.001 | -60.0057 | -0.000926 | -0.0001995 |  |  | 32 | 4.0315 | [7 empty] | 50
2023-02-16 |  | CHIPs retro (after RM optimization) | 20.521 | 23.806 | 21.691 | 23.884 | 24.63 | 23.23 | -2147.48 | -65.001 | -60.0057 | -0.000926 | -0.0001995 |  |  | 8 | 4.7624 | [8 empty] | 10:03 | 302.2 | 300.4 | 300.1 | 300.9 | 302.4 | 299.5 | 300.1 | disabled
2023-02-16 |  | PATB mirror | 20.521 | 23.806 | 21.691 | 23.884 | 24.13 | 23.77 | 101.502 | 125.708 | -0.0006 | 0.108 | 0.05334 | 0,20 |  |  |  | 31526 |  | 0.0017 | 0.0021
2023-02-16 |  | PATB LED pri | 20.521 | 23.806 | 21.691 | 23.884 | 24.13 | 23.77 | 88.2305 | 108.3 | -0.0006 | 0.1081 | 0.0534 | 1132 |  |  |  | 32553 |  |  |  | 801.6 | 601.1
2023-02-16 |  | PATB Retro pri | 20.521 | 23.806 | 21.691 | 23.884 | 24.13 | 23.77 | 67.6974 | 103.398 | -0.0008 | 0.1081 | 0.0534 |  |  | 20 | 3.459
2023-02-16 |  | PATB Retro pri (@ sMATF pinhole) | 20.521 | 23.806 | 21.691 | 23.884 | 24.13 | 23.77 | 100.414 | 100.4 | -0.0003 | 0.1081 | 0.0534 | 1140 |  |  |  | 34049 |  |  |  | 807 | 601.5
Date |  |  | RM1 |  | RM2 |  | RM3 |  | 5DOFx (mm) | 5DOFy (mm) | 5DOFz (mm) | 5DOF Rx (deg) | 5DOF Ry (deg) | AC Focus | Cam exp(ms) | Lidar Sig | Range (m) | .fits | .bmp | AC AZ | AC EL | X centroid | Y centroid
2023-02-18 |  | CHIPs retro | 20.521 | 23.806 | 21.691 | 23.884 | 24.63 | 23.23 | -2147.48 | -65 | -60.0065 | -0.0003 | 8.8e-05 |  |  | 10 | 4.7624
2023-02-18 |  | sMATF retro | 20.521 | 23.806 | 21.691 | 23.884 | 24.13 | 23.77 | -2147.48 | -65 | -60.0065 | -0.0003 | 8.8e-05 |  |  | 25 | 4.0315
2023-02-18 |  | sMATF mirror (after RM optimization) | 20.517 | 23.802 | 21.691 | 23.884 | 24.13 | 23.77 | -2147.48 | -65 | -60.0065 | -0.0003 | 8.8e-05 | 0,20 |  |  |  | 32425 |  | 0.0027 | 0.0028
2023-02-18 |  | sMATF LED (after RM optimization) | 20.517 | 23.802 | 21.691 | 23.884 | 24.13 | 23.77 | -2147.48 | -65 | -60.0065 | -0.0003 | 8.8e-05 | 1,76 |  |  |  | 32604 |  |  |  | 807.7 | 602.6
2023-02-18 |  | sMATF retro (after RM optimization) | 20.517 | 23.802 | 21.691 | 23.884 | 24.13 | 23.77 | -2147.48 | -65 | -60.0065 | -0.0003 | 8.8e-05 |  |  | 35 | 4.0315
2023-02-18 |  | CHIPs retro (after RM optimization) | 20.521 | 23.806 | 21.691 | 23.884 | 24.63 | 23.23 | -2147.48 | -65 | -60.0065 | -0.0003 | 8.8e-05 |  |  | 11 | 4.7624
2023-02-18 |  | PATB mirror | 20.517 | 23.802 | 21.691 | 23.884 | 24.13 | 23.77 | 101.501 | 125.705 | -0.00014 | 0.1087 | 0.063417 | 0,20 |  |  |  | 43911 |  | 0.0027 | 0.0049
2023-02-18 |  | PATB LED pri | 20.517 | 23.802 | 21.691 | 23.884 | 24.13 | 23.77 | 88.5758 | 108.3 | -0.0002 | 0.1088 | 0.063452 | 1132 |  |  |  | 44501 |  |  |  | 807.6 | 601.8
2023-02-18 |  | PATB Retro pri | 20.517 | 23.802 | 21.691 | 23.884 | 24.13 | 23.77 | 67.6969 | 103.398 | -0.00047 | 0.10883 | 0.063471 |  |  | 15 | 3.4591
2023-02-18 |  | PATB Retro pri (@ sMATF pinhole) | 20.517 | 23.802 | 21.691 | 23.884 | 24.13 | 23.77 | 100.717 | 100.451 | 0 | 0.10887 | 0.06349 | 1140 |  |  |  | 45954 |  |  |  | 808 | 602.5
POST PR5
2023-02-19 |  | CHIPs retro | 20.517 | 23.802 | 21.691 | 23.884 | 24.63 | 23.23 | -2147.480(+250) | -65 | -60 | 0 | 0 |  |  | 11 | 4.7624 | [8 empty] | 13:17 | 304.6 | 300.9 | 300.8 | 302.1 | 302.5 | 300.3 | 301.8 | disabled
2023-02-19 |  | sMATF retro | 20.517 | 23.802 | 21.691 | 23.884 | 24.13 | 23.77 | -2147.480(+250) | -65 | -60 | 0 | 0 |  |  | 24 | 4.0315
2023-02-19 |  | sMATF mirror (after RM optimization) | 20.5205 | 23.8045 | 21.691 | 23.884 | 24.13 | 23.77 | -2147.480(+250) | -65 | -60 | 0 | 0 | 0,17 |  |  |  | 34019 |  | 0.0013 | 0.0025
2023-02-19 |  | sMATF LED (after RM optimization) | 20.5205 | 23.8045 | 21.691 | 23.884 | 24.13 | 23.77 | -2147.480(+250) | -65 | -60 | 0 | 0 | 1,76 |  |  |  | 33843 |  |  |  | 807.7 | 601.7
2023-02-19 |  | sMATF retro (after RM optimization) | 20.5205 | 23.8045 | 21.691 | 23.884 | 24.13 | 23.77 | -2147.480(+250) | -65 | -60 | 0 | 0 |  |  | 24 | 4.0315
2023-02-19 |  | CHIPs retro (after RM optimization) | 20.5205 | 23.8045 | 21.691 | 23.884 | 24.63 | 23.23 | -2147.480(+250) | -65 | -60 | 0 | 0 |  |  | 12 | 4.7624
2023-02-19 |  | PATB mirror | 20.5205 | 23.8045 | 21.691 | 23.884 | 24.13 | 23.77 | 101.501 | 125.71 | -0.0001 | 0.1087 | 0.06356 | 0,20 |  |  |  | 35923 |  | -0.0022 | 0.0084 |  |  |  |  | 13:56 | 304.7 | 301 | 300.8 | 302.4 | 304.1 | 300.4 | 302.4 | disabled
2023-02-19 |  | PATB LED pri | 20.5205 | 23.8045 | 21.691 | 23.884 | 24.13 | 23.77 | 88.3713 | 108.246 | -0.0009 | 0.1088 | 0.0636 | 1131 |  |  |  | 40651 |  |  |  | 807.7 | 601.6
2023-02-19 |  | PATB Retro pri | 20.5205 | 23.8045 | 21.691 | 23.884 | 24.13 | 23.77 | 67.6964 | 103.398 | -0.0011 | 0.1088 | 0.0636 |  |  | 16 | 3.459
2023-02-19 |  | PATB Retro pri (@ sMATF pinhole) | 20.5205 | 23.8045 | 21.691 | 23.884 | 24.13 | 23.77 | 100.478 | 103.3515 | -0.0057 | 0.1089 | 0.0637 | 1132 |  |  |  | 43246 |  |  |  | 807 | 601 | 128.2 |  | 14:37 | 304.9 | 301 | 300.9 | 302.4 | 306.5 | 300.5 | 302.8 | disabled
Date |  |  | RM1 |  | RM2 |  | RM3 |  | 5DOFx (mm) | 5DOFy (mm) | 5DOFz (mm) | 5DOF Rx (deg) | 5DOF Ry (deg) | AC Focus | Cam exp(ms) | Lidar Sig | Range (m) | .fits | .bmp | AC AZ | AC EL | X centroid | Y centroid
2023-02-21 |  | PATB LED pri | 20.5205 | 23.8045 | 21.691 | 23.884 | 24.13 | 23.77 | 88.373 | 108.247 |  | 0.108 | 0.063 | 1132 | [7 empty] | 809.7 | 601.8
|  | Jog AC focus to 126 and + to 129, PSF not optimal by eye | 20.5205 | 23.8045 | 21.691 | 23.884 | 24.13 | 23.77 | [5 empty] | 1129 |  |  |  | 4953 |  |  |  | 808.9 | 602.1
|  |  | 20.5205 | 23.8045 | 21.691 | 23.884 | 24.13 | 23.77 | [5 empty] | 1130 |  |  |  | 5118 |  |  |  | 809 | 602.6
|  |  | 20.5205 | 23.8045 | 21.691 | 23.884 | 24.13 | 23.77 | [5 empty] | 1131 |  |  |  | 5229 |  |  |  | 809 | 602
|  |  | 20.5205 | 23.8045 | 21.691 | 23.884 | 24.13 | 23.77 | [5 empty] | 1132 |  |  |  | 5514 |  |  |  | 809 | 602
|  |  | 20.5205 | 23.8045 | 21.691 | 23.884 | 24.13 | 23.77 | [5 empty] | 1133 |  |  |  | 5709 |  |  |  | 809 | 602
|  | PSF degredated by eye | 20.5205 | 23.8045 | 21.691 | 23.884 | 24.13 | 23.77 | [5 empty] | 1134 |  |  |  | 5851 |  |  |  | 809 | 602
|  | Optimal PSF by eye | 20.5205 | 23.8045 | 21.691 | 23.884 | 24.13 | 23.77 | [5 empty] | 1130 |  |  |  | 10652 |  |  |  | 808.1 | 602.6
|  | Home Rm1A, and replace | 20.5205 | 23.8045 | 21.691 | 23.884 | 24.13 | 23.77 | [5 empty] | 1130 | [7 empty] | 806.8 | 608.8
|  | Tune Rm1A, B, wait a few minutes and retune 1B | 20.528 | 23.809 | 21.691 | 23.884 | 24.13 | 23.77 | [5 empty] | 1130 | [7 empty] | 808 | 601.7
|  | Home and Replace RM1 | 20.528 | 23.809 | 21.691 | 23.884 | 24.13 | 23.77 | [13 empty] | 900.9 | 729.7
|  | PATB pri Retro at center position, image with AC LED illumination | 20.64 | 23.74 | 21.691 | 23.884 |  |  | 100.48 | 100.352 |  | 0.108 | 0.064 | 1130 |  |  |  | 21340
|  | PATB pri Retro at center position, image with LIDAR illumination | 20.64 | 23.74 | 21.691 | 23.884 | 24.35 | 23.42 | [5 empty] | 1149 |  |  |  | 31221 | [5 empty] | 125
|  | PATB pri Retro at center position, image with AC LED illumination | 20.64 | 23.74 | 21.691 | 23.884 | [7 empty] | 1127 |  |  |  | 32025 | [5 empty] | 130
|  | PATB LED pri | 20.5205 | 23.804 | 21.691 | 23.884 | [7 empty] | 1133 | [7 empty] | 807.7 | 602
|  | PATB pri Retro at center position, image with LIDAR illumination | 20.5205 | 23.804 | 21.691 | 23.884 | [7 empty] | 1121 |  |  |  | 40226
|  | PATB pri Retro at center position, image with AC LED illumination | 20.5205 | 23.804 | 21.691 | 23.884 | 24.4 | 23.49 | [5 empty] | 1143 |  |  |  | 42830 |  |  |  | 805 | 600
|  | PATB LED pri position, Retro is also captured in image | [6 empty] | 88.375 | 108.245 |  | 0.108 | 0.0637 |  |  |  |  | 51921 |  |  |  | 807.3 | 601.8
Date |  |  | RM1 |  | RM2 |  | RM3 |  | 5DOFx (mm) | 5DOFy (mm) | 5DOFz (mm) | 5DOF Rx (deg) | 5DOF Ry (deg) | AC Focus | Cam exp(ms) | Lidar Sig | Range (m) | .fits | .bmp | AC AZ | AC EL | X centroid | Y centroid
2023-02-22 |  | CHIPs retro | 20.5205 | 23.804 | 21.691 | 23.884 | 24.63 | 23.23 | -2147.480(+250) | -65 | -60 | 0 | 0 |  |  | 13 | 4.7624
2023-02-22 |  | sMATF retro | 20.5205 | 23.804 | 21.691 | 23.884 | 24.14 | 23.77 | -2147.480(+250) | -65 | -60 | 0 | 0 |  |  | 38 | 4.0315
2023-02-22 |  | sMATF mirror (after RM optimization) | 20.472 | 23.805 | 21.6928 | 23.8405 | 24.14 | 23.77 | -2147.480(+250) | -65 | -60 | 0 | 0 | 0,14 |  | 18 | 4.021 | 61413 |  | -0.0015 | 0.0028 |  |  |  | 0.54943
2023-02-22 |  | sMATF LED (after RM optimization) | 20.472 | 23.805 | 21.6928 | 23.8405 | 24.14 | 23.77 | -2147.480(+250) | -65 | -60 | 0 | 0 | 1,70 |  |  |  | 61639 |  |  |  | 806.7 | 603.6
2023-02-22 |  | sMATF retro (after RM optimization) | 20.472 | 23.805 | 21.6928 | 23.8405 | 24.14 | 23.77 | -2147.480(+250) | -65 | -60 | 0 | 0 |  |  | 24 | 4.0315
2023-02-22 |  | CHIPs retro (after RM optimization) | 20.472 | 23.805 | 21.6928 | 23.8405 | 24.63 | 23.23 | -2147.480(+250) | -65 | -60 | 0 | 0 |  |  | 13 | 4.7624
2023-02-22 |  | PATB mirror | 20.472 | 23.805 | 21.6928 | 23.8405 | 24.44 | 23.475 | 101.502 | 125.708 | 0 | 0.10837 | 0.06332 | 0,20 |  | 41 | 3.4716 | 64906 |  | -0.0049 | 0.0067
2023-02-22 |  | PATB LED pri | 20.472 | 23.805 | 21.6928 | 23.8405 | 24.44 | 23.475 | 88.3707 | 108.245 | 0 | 0.1085 | 0.0633 | 1128 |  |  |  | 70403 |  |  |  | 806.8 | 604.6
2023-02-22 |  | PATB Retro pri (@ sMATF pinhole) | 20.472 | 23.805 | 21.6928 | 23.8405 | 24.395 | 23.475 | 100.479 | 100.351 | 0.0002 | 0.10854 | 0.0634 | 1142 |  | 31 | 3.4582 | 71815 |  |  |  | 806 | 602
2023-02-22 |  | PATB Retro pri | 20.472 | 23.805 | 21.6928 | 23.8405 | 24.14 | 23.77 | 67.6949 | 103.398 | -0.0001 | 0.10855 | 0.06341 |  |  | 18 | 3.4591
Shim happened here
Date |  |  | RM1 |  | RM2 |  | RM3 |  | 5DOFx (mm) | 5DOFy (mm) | 5DOFz (mm) | 5DOF Rx (deg) | 5DOF Ry (deg) | AC Focus | Cam exp(ms) | Lidar Sig | Range (m) | .fits | .bmp | AC AZ | AC EL | X centroid | Y centroid
2023-02-24 |  | CHIPs retro | 20.472 | 23.805 | 21.6928 | 23.8405 | 24.63 | 23.23 | -2147.480(+250) | -65.001 | -60.004 | -0.0004 | -0.0003 |  |  | 7 | 4.7617
2023-02-24 |  | sMATF retro | 20.472 | 23.805 | 21.6928 | 23.8405 | 24.14 | 23.77 | -2147.480(+250) | -65 | -60 | 0 | 0 |  |  | 31 | 4.0309
2023-02-24 |  | sMATF mirror (after RM optimization) | 21.367 | 23.197 | 20.8343 | 24.6731 | 24.14 | 23.77 | -2147.480(+250) | -65 | -60 | 0 | 0 | 0,21 |  |  |  | 43426 |  | -0.0015 | 0.0053
2023-02-24 |  | sMATF LED (after RM optimization) | 21.367 | 23.197 | 20.8343 | 24.6731 | 24.14 | 23.77 | -2147.480(+250) | -65 | -60 | 0 | 0 | 1,70 |  |  |  | 43554 |  |  |  | 807.2 | 601.9
2023-02-24 |  | Mirror Range (after RM optimization) Max Signal but bad StDev | 21.367 | 23.197 | 20.8343 | 24.6731 | 24.44 | 23.49 | -2147.480(+250) | -65 | -60 | 0 | 0 |  |  | 90 | 4.0212 | [7 empty] | 0.000204
2023-02-24 |  | CHIPs retro (after RM optimization) | 21.367 | 23.197 | 20.8343 | 24.6731 | 24.633 | 23.23 | -2147.480(+250) | -65 | -60 | 0 | 0 |  |  | 7 | 4.7619
2023-02-24 |  | sMATF retro (after RM optimization) | 21.367 | 23.197 | 20.8343 | 24.6731 | 24.14 | 23.77 | -2147.480(+250) | -65 | -60 | 0 | 0 |  |  | 34 | 4.031
2023-02-24 |  | PATA mirror | 21.367 | 23.197 | 20.8343 | 24.6731 | 24.14 | 23.77 | -76.1026 | -0.0025 | 125.704 | -0.0093 | 0.0501 | 0,20 |  |  |  | 54757 |  | 0.0006 | 0.0011
2023-02-24 |  | PATA Retro pri | 21.367 | 23.197 | 20.8343 | 24.6731 | 24.14 | 23.77 | -105.496 | 102.7 | 0.0007 | -0.0092 | 0.0501 |  |  | 13 | 3.459
2023-02-24 |  | PATA LED pri | 21.367 | 23.197 | 20.8343 | 24.6731 | 24.14 | 23.77 | -60.4297 | 107.399 | -0.0031 | -0.0092 | 0.0501 | 1135 |  |  |  | 60601 |  |  |  | 807.7 | 601.7
Date |  |  | RM1 |  | RM2 |  | RM3 |  | 5DOFx (mm) | 5DOFy (mm) | 5DOFz (mm) | 5DOF Rx (deg) | 5DOF Ry (deg) | AC Focus | Cam exp(ms) | Lidar Sig | Range (m) | .fits | .bmp | AC AZ | AC EL | X centroid | Y centroid
2023-02-25 |  | CHIPs retro | 21.367 | 23.197 | 20.8343 | 24.6731 | 24.63 | 23.23 | -2147.480(+250) | -65.001 | -60.004 | -0.0004 | -0.0003 |  |  | 6 | 4.7619
2023-02-25 |  | sMATF retro, some high StDev points up to 11um | 21.367 | 23.197 | 20.8343 | 24.6731 | 24.14 | 23.77 | -2147.480(+250) | -65.001 | -60.004 | -0.0004 | -0.0003 |  |  | 32 | 4.031 | [7 empty] | -4.1e-05
2023-02-25 |  | sMATF mirror | 21.367 | 23.197 | 20.8343 | 24.6731 | 24.14 | 23.77 | -2147.480(+250) | -65.001 | -60.004 | -0.0004 | -0.0003 | [6 empty] | -0.0015 | 0.0084 | 809.5 | 602.6
2023-02-25 |  | sMATF mirror (after RM optimization) | 21.3915 | 23.2385 | 20.8863 | 24.6866 | 24.14 | 23.77 | -2147.480(+250) | -65.001 | -60.004 | -0.0004 | -0.0003 | 0,15 |  |  |  | 12159 |  | -0.0028 | 0.0032
2023-02-25 |  | sMATF LED (after RM optimization) | 21.3915 | 23.2385 | 20.8863 | 24.6866 | 24.14 | 23.77 | -2147.480(+250) | -65.001 | -60.004 | -0.0004 | -0.0003 | 1,67 |  |  |  | 11153 |  |  |  | 807.5 | 601.6
2023-02-25 |  | Mirror Range (after RM optimization) all bad StDev except first point | 21.3915 | 23.2385 | 20.8863 | 24.6866 | 24.442 | 23.49 | -2147.480(+250) | -65.001 | -60.004 | -0.0004 | -0.0003 |  |  | 22 | 4.023
2023-02-25 |  | CHIPs retro (after RM optimization) | 21.3915 | 23.2385 | 20.8863 | 24.6866 | 24.63 | 23.23 | [7 empty] | 6 | 4.7619 | [7 empty] | 3.6e-05
2023-02-25 |  | sMATF retro (after RM optimization) | 21.3915 | 23.2385 | 20.8863 | 24.6866 | 24.14 | 23.77 | [8 empty] | 4.0311
2023-02-25 |  | Mirror Range (after RM optimization) all bad StDev except first point | 21.3915 | 23.2385 | 20.8863 | 24.6866 | 24.442 | 23.49 | -2147.480(+250) | -65.001 | -60.004 | -0.0004 | -0.0003 |  |  | 43 | 4.0224 | [7 empty] | -0.000589
2023-02-25 |  | PATB mirror | 21.3915 | 23.2385 | 20.8863 | 24.6866 | 24.442 | 23.49 | 101.504 | 125.712 | 0 | 0.0132 | 0.0621 | 0,15 |  |  |  | 25329 |  | -0.0025 | 0.0035
2023-02-25 |  | Range PATB mirror, all bad StDev except first point | 21.3915 | 23.2385 | 20.8863 | 24.6866 | 24.442 | 23.49 | 101.504 | 125.712 | 0 | 0.0132 | 0.0621 | 0,15 |  | 138 | 3.4737
2023-02-25 |  | PATB LED pri | 21.3915 | 23.2385 | 20.8863 | 24.6866 | 24.442 | 23.49 | 88.4729 | 107.436 | 0 | 0.01327 | 0.0622 | 1135 | [7 empty] | 806.9 | 601.8
2023-02-25 |  | PATB LED pri | 21.3915 | 23.2385 | 20.8863 | 24.6866 | 24.442 | 23.49 | 97.468 | 107.436 | 0 | 0.01327 | 0.0622 | 1135 | [7 empty] | 964.4 | 497.6
2023-02-25 |  | PATB LED pri | 21.3915 | 23.2385 | 20.8863 | 24.6866 | 24.442 | 23.49 | 79.4819 | 107.436 | 0 | 0.01327 | 0.0622 | 1135 | [7 empty] | 649.9 | 705.9
2023-02-25 |  | PATB LED pri | 21.3915 | 23.2385 | 20.8863 | 24.6866 | 24.442 | 23.49 | 79.4821 | 116.436 | 0 | 0.01327 | 0.0622 | 1135 | [7 empty] | 546.6 | 551.4
2023-02-25 |  | PATB LED pri | 21.3915 | 23.2385 | 20.8863 | 24.6866 | 24.442 | 23.49 | 79.4821 | 98.436 | 0 | 0.01327 | 0.0622 | 1135 | [7 empty] | 751.97 | 860.481
2023-02-25 |  | PATB LED pri | 21.3915 | 23.2385 | 20.8863 | 24.6866 | 24.442 | 23.49 | 88.468 | 98.435 | 0 | 0.01327 | 0.0622 | 1135 | [7 empty] | 909.1 | 755.8
2023-02-25 |  | PATB LED pri | 21.3915 | 23.2385 | 20.8863 | 24.6866 | 24.442 | 23.49 | 88.469 | 116.436 | 0 | 0.01327 | 0.0622 | 1135 | [7 empty] | 704.8 | 447.4
2023-02-25 |  | PATB Retro pri, pinhole location | 21.3915 | 23.2385 | 20.8863 | 24.6866 |  |  | 102.48 | 101.35 | 0 | 0.01327 | 0.0622 | 1142 |  |  |  | 41640 |  |  |  | 822 | 549
2023-02-25 |  | PATB Retro pri, Apex and LED Pri image | 21.3915 | 23.2385 | 20.8863 | 24.6866 | 24.442 | 23.49 | 66.6936 | 102.398 | 0 | 0.01327 | 0.0622 | 1123 |  |  | 3.4589 | 43748 |  |  |  | 177 | 942
2023-02-25 |  | PATB Retro pri, Range | 21.3915 | 23.2385 | 20.8863 | 24.6866 | 24.442 | 23.49 | 66.6936 | 102.398 | 0 | 0.01327 | 0.0622 | 1123 |  |  | 3.4589 | 44126
Date |  |  | RM1 |  | RM2 |  | RM3 |  | 5DOFx (mm) | 5DOFy (mm) | 5DOFz (mm) | 5DOF Rx (deg) | 5DOF Ry (deg) | AC Focus | Cam exp(ms) | Lidar Sig | Range (m) | .fits | .bmp | AC AZ | AC EL | X centroid | Y centroid
2023-02-27 |  | CHIPs retro | 21.3915 | 23.2385 | 20.8863 | 24.6866 | 24.64 | 23.24 | -2147.480(+250) | -65.001 | -60.004 | -0.0004 | -0.0003 |  |  | 7 | 4.7619
2023-02-27 |  | sMATF retro | 21.3915 | 23.2385 | 20.8863 | 24.6866 | 24.14 | 23.77 | -2147.480(+250) | -65 | -60 | 0 | 0 |  |  | 35 | 4.031
2023-02-27 |  | sMATF mirror (after RM optimization) | 21.3825 | 23.2436 | 20.913 | 24.6903 | 24.14 | 23.77 | -2147.480(+250) | -65 | -60 | 0 | 0 | 0,15 |  |  |  | 22633 |  | -0.0035 | 0.0035
2023-02-27 |  | sMATF LED (after RM optimization) | 21.3825 | 23.2436 | 20.913 | 24.6903 | 24.14 | 23.77 | -2147.480(+250) | -65 | -60 | 0 | 0 | 1,67 |  |  |  | 22435 |  |  |  | 807.8 | 601.6
2023-02-27 |  | CHIPs retro (after RM optimization) | 21.3915 | 23.2385 | 20.8863 | 24.6866 | 24.64 | 23.24 | -2147.480(+250) | -65 | -60 | 0 | 0 |  |  | 7 | 4.7619
2023-02-27 |  | sMATF retro (after RM optimization) | 21.367 | 23.197 | 20.8343 | 24.6731 | 24.14 | 23.77 | -2147.480(+250) | -65 | -60 | 0 | 0 |  |  | 43 | 4.031
2023-02-27 |  | PATB mirror | 21.367 | 23.197 | 20.8343 | 24.6731 | 24.14 | 23.77 | 101.503 | 125.709 | 0.0016 | 0.01276 | 0.06193 | 0,15 |  |  |  | 25239 |  | -0.0063 | 0.0035
2023-02-27 |  | PATB retro pri | 21.367 | 23.197 | 20.8343 | 24.6731 | 24.14 | 23.77 | 66.6936 | 102.3984 | 0 | 0.01276 | 0.06193 |  |  | 17 | 3.4589
2023-02-27 |  | PATB LED pri | 21.367 | 23.197 | 20.8343 | 24.6731 | 24.14 | 23.77 | 88.3854 | 107.388 | 0.0013 | 0.01276 | 0.06193 | 1122 |  |  |  | 30738 |  |  |  | 807.6 | 601.7
Date |  |  | RM1 |  | RM2 |  | RM3 |  | 5DOFx (mm) | 5DOFy (mm) | 5DOFz (mm) | 5DOF Rx (deg) | 5DOF Ry (deg) | AC Focus | Cam exp(ms) | Lidar Sig | Range (m) | .fits | .bmp | AC AZ | AC EL | X centroid | Y centroid
2023-03-14 |  | sMATF mirror | 21.3915 | 23.2385 | 20.8863 | 24.6866 |  |  | -2147.480(+250) | -65.001 | -60.004 | [8 empty] | 0.06 | 0.015
2023-03-14 |  | sMATF LED | 21.3915 | 23.2385 | 20.8863 | 24.6866 |  |  | -2147.480(+250) | -65 | -60
2023-03-14 |  | Range sMATF mirror |  |  |  |  | 24.443 | 23.491 | 101.503 | 125.709 | 0.0016 |  |  |  |  | 14 | 4.0214
Date |  |  | RM1 |  | RM2 |  | RM3 |  | 5DOFx (mm) | 5DOFy (mm) | 5DOFz (mm) | 5DOF Rx (deg) | 5DOF Ry (deg) | AC Focus | Cam exp(ms) | Lidar Sig | Range (m) | .fits | .bmp | AC AZ | AC EL | X centroid | Y centroid
2023-03-15 |  | sMATF retro | 21.3915 | 23.2385 | 20.8863 | 24.6866 | 24.115 | 23.705 | -2147.480(+250) | -65 | -60 |  |  |  |  | 32 | 4.0319
2023-03-15 |  | Range sMATF mirror | 21.3915 | 23.2385 | 20.8863 | 24.6866 | 24.443 | 23.47 | 101.503 | 125.709 | 0.0016 |  |  |  |  | 7 | 4.0214
2023-03-15 |  | Range sMATF mirror | 21.3915 | 23.2385 | 20.8863 | 24.6866 | 24.444 | 23.47 | 101.503 | 125.709 | 0.0016 |  |  |  |  | 5 | 4.0215
2023-03-15 |  | Range sMATF mirror | 21.3915 | 23.2385 | 20.8863 | 24.6866 | 24.438 | 23.47 | [7 empty] | 8 | 4.0215
2023-03-15 |  | Range sMATF mirror | 21.3915 | 23.2385 | 20.8863 | 24.6866 | 24.439 | 23.47 | [7 empty] | 10 | 4.0215
2023-03-15 |  | Range sMATF mirror | 21.3915 | 23.2385 | 20.8863 | 24.6866 | 24.44 | 23.47 | [7 empty] | 9 | 4.0215
2023-03-15 |  | Range sMATF mirror | 21.3915 | 23.2385 | 20.8863 | 24.6866 | 24.441 | 23.47 | [7 empty] | 11 | 4.0216
2023-03-15 |  | Range sMATF mirror | 21.3915 | 23.2385 | 20.8863 | 24.6866 | 24.442 | 23.47 | [7 empty] | 10 | 4.0216
2023-03-15 |  | Range sMATF mirror | 21.3915 | 23.2385 | 20.8863 | 24.6866 | 24.443 | 23.47 | [7 empty] | 9 | 4.0215
2023-03-15 |  | Range sMATF mirror | 21.3915 | 23.2385 | 20.8863 | 24.6866 | 24.444 | 23.47 | [7 empty] | 7 | 4.0215
2023-03-15 |  | Range sMATF mirror | 21.3915 | 23.2385 | 20.8863 | 24.6866 | 24.445 | 23.47 | [7 empty] | 7 | 4.0215 | [8 empty] | 16:37 | 311.4 | 307.8 | 307.6 | 309.2 | 309.8 | 307 | 309.2 | disabled
2023-03-15 |  | sMATF mirror before nulling | 21.3915 | 23.2385 | 20.8863 | 24.6866 | 24.445 | 23.47 | [11 empty] | 0.057 | 0.0094
2023-03-15 |  | sMATF mirror before nulling | 21.3915 | 23.2385 | 20.8863 | 24.6866 | 24.445 | 23.47 | [13 empty] | 772.5 | 579.7
2023-03-15 |  | sMATF mirror Rm1, 2 Nulling | 20.441 | 23.892 | 21.7805 | 23.8222 | 24.445 | 23.47 | [5 empty] | 1,75 |  |  |  | 81207 |  |  |  | 807.6 | 602.5 |  |  | 5:46 | 311.8 | 308.3 | 308 | 309.4 | 309.4 | 307.5 | 309.3
2023-03-15 |  | sMATF mirror Rm1, 2 Nulled 6:20pm | 20.441 | 23.892 | 21.7805 | 23.8222 | 24.445 | 23.47 | [5 empty] | 0,20 |  |  |  | 81933 |  | -0.0001 | -0.0034
2023-03-15 |  | sMATF mirror Rm1, 2 Nulled 6:27pm | 20.441 | 23.892 | 21.7805 | 23.8222 | 24.445 | 23.47 | [5 empty] | 0,21 |  |  |  | 82751 |  | 0.0003 | 0.0001
2023-03-15 |  | sMATF mirror Rm1, 2 Nulled 6:38pm | 20.441 | 23.892 | 21.7805 | 23.8222 | 24.445 | 23.47 | [5 empty] | 1,75 |  |  |  | 83821 |  |  |  | 811.6 | 601.6 |  |  | 18:39 | 312 | 308.6 | 308.2 | 309.7 | 313.9 | 307.8 | 309.9
2023-03-15 |  | sMATF mirror Rm1, 2 Nulled 6:46pm | 20.441 | 23.892 | 21.7805 | 23.8222 | 24.445 | 23.47 | [5 empty] | 1,76 |  |  |  | 84623 |  |  |  | 811.6 | 601.6
2023-03-15 |  | sMATF mirror Rm1, 2 Nulled 6:49pm | 20.441 | 23.892 | 21.7805 | 23.8222 | 24.445 | 23.47 | [5 empty] | 0,21 |  |  |  | 84952 |  | 0.0006 | 0.0011
2023-03-15 |  | sMATF mirror Rm1, 2 Nulled 7:21pm | 20.441 | 23.892 | 21.7805 | 23.8222 | 24.445 | 23.47 | [5 empty] | 0,22 |  |  |  | 92110 |  | -0.0001 | 0.0015
2023-03-15 |  | sMATF mirror Rm1, 2 Nulled 7:23pm | 20.441 | 23.892 | 21.7805 | 23.8222 | 24.445 | 23.47 | [5 empty] | 1,76 |  |  |  | 92356 |  |  |  | 813.5 | 600.6
Date |  |  | RM1 |  | RM2 |  | RM3 |  | 5DOFx (mm) | 5DOFy (mm) | 5DOFz (mm) | 5DOF Rx (deg) | 5DOF Ry (deg) | AC Focus | Cam exp(ms) | Lidar Sig | Range (m) | .fits | .bmp | AC AZ | AC EL | X centroid | Y centroid
2023-03-16 |  | sMATF mirror, record drift | 20.441 | 23.892 | 21.7805 | 23.8222 | 24.445 | 23.47 | [5 empty] | 0,22 |  |  |  | 84402 |  | -0.0022 | 0.0022 |  |  |  |  | ~8:45am
2023-03-16 |  | sMATF mirror, record drift | 20.441 | 23.892 | 21.7805 | 23.8222 | 24.445 | 23.47 | [5 empty] | 1,76 |  |  |  | 85706 |  |  |  | 810.7 | 597 |  |  | ~9:00am
2023-03-16 |  | CHIPs retro | 20.441 | 23.892 | 21.7805 | 23.8222 | 24.64 | 23.24 | [16 empty] | couldnt get signal
2023-03-16 |  | sMATF retro | 20.441 | 23.892 | 21.7805 | 23.8222 | 24.13 | 23.72 | [7 empty] | 31 | 4.0316
2023-03-16 |  | sMATF mirror, record drift | 20.441 | 23.892 | 21.7805 | 23.8222 | [7 empty] | 0,22 |  |  |  | 1542 |  | -0.0022 | 0.0022 |  |  |  |  | ~10:15am
2023-03-16 |  | sMATF mirror, record drift | 20.441 | 23.892 | 21.7805 | 23.8222 | [7 empty] | 1,76 |  |  |  | 1325 |  |  |  | 812 | 605.7 |  |  | ~10:15am
2023-03-16 |  | sMATF mirror, record drift | 20.441 | 23.892 | 21.7805 | 23.8222 | 24.64 | 23.24 | [5 empty] | 1,76 |  |  |  | 11606 |  |  |  | 812.9 | 604.7 |  |  | ~11:15 | 308 | 305 | 304.8 | 305.9 | 305.8 | 304.2 | 305.5
2023-03-16 |  | sMATF mirror, record drift | 20.441 | 23.892 | 21.7805 | 23.8222 | 24.13 | 23.72 | [5 empty] | 1,76 |  |  |  | 11838 |  |  |  | 812.9 | 604.7 |  |  | ~11:20
2023-03-16 |  | sMATF mirror, record drift | 20.441 | 23.892 | 21.7805 | 23.8222 | 24.13 | 23.72 | [5 empty] | 1,76 |  |  |  | 13209 |  |  |  | 812.8 | 604.7 |  |  | ~11:30 | 308.1 | 305.2 | 305 | 306 | 307.7 | 304.3 | 305.7
2023-03-16 |  | sMATF mirror, record drift | 20.441 | 23.892 | 21.7805 | 23.8222 | 24.13 | 23.72 | [5 empty] | 1,76 |  |  |  | 13536 |  |  |  | 812.8 | 604.8 |  |  | ~11:35
2023-03-16 |  | sMATF mirror, record drift | 20.441 | 23.892 | 21.7805 | 23.8222 | 24.13 | 23.72 | [5 empty] | 1,76 |  |  |  | 15150 |  |  |  | 811.9 | 604.8 |  | just turned off RM controller | ~11:51 | 309.6 | 305.3 | 305.1 | 306.2 | 309.6 | 304.5 | 306
2023-03-16 |  | sMATF mirror, record drift | 20.441 | 23.892 | 21.7805 | 23.8222 | 24.13 | 23.72 | [5 empty] | 1,76 |  |  |  | 20224 |  |  |  | 811.7 | 602.6 |  |  | ~12:03 | 308.4 | 305.4 | 305.2 | 306.2 | 310 | 304.5 | 306.3
2023-03-16 |  | sMATF mirror, record drift | 20.441 | 23.892 | 21.7805 | 23.8222 | 24.13 | 23.72 | [5 empty] | 1,76 |  |  |  | 20956 |  |  |  | 810.9 | 600.6 |  |  | ~12:11 | 308.3 | 305.5 | 305.2 | 306.1 | 310.2 | 304.5 | 306.4
2023-03-16 |  | sMATF mirror, record drift | 20.441 | 23.892 | 21.7805 | 23.8222 | 24.13 | 23.72 | [5 empty] | 1,76 |  |  |  | 21932 |  |  |  | 810.7 | 598.7 |  |  | ~12:20 | 308.1 | 305.5 | 305.3 | 306.1 | 310.2 | 304.5 | 306.4
2023-03-16 |  | sMATF mirror, record drift | 20.441 | 23.892 | 21.7805 | 23.8222 | 24.13 | 23.72 | [5 empty] | 1,76 |  |  |  | 22925 |  |  |  | 810.5 | 597.6 |  |  | ~12:30 | 308 | 305.5 | 305.3 | 306 | 310.3 | 304.5 | 306.4
2023-03-16 |  | sMATF mirror, record drift | 20.441 | 23.892 | 21.7805 | 23.8222 | 24.13 | 23.72 | [5 empty] | 1,76 |  |  |  | 24003 |  |  |  | 809.8 | 596.7 |  | camera off after this measurement to cool down | ~12:40 | 307.8 | 305.5 | 305.4 | 306 | 310.3 | 304.6 | 306.4
2023-03-16 |  | sMATF mirror, record drift | 20.441 | 23.892 | 21.7805 | 23.8222 | 24.13 | 23.72 | [5 empty] | 1,76 |  |  |  | 35344 |  |  |  | 809.9 | 596.6 |  | camera back on for more measurements | ~13:50 | 307.2 | 305.6 | 305.4 | 305.7 | 305.2 | 304.6 | 305.5
2023-03-16 |  | sMATF mirror, record drift | 20.441 | 23.892 | 21.7805 | 23.8222 | 24.13 | 23.72 | [5 empty] | 1,76 |  |  |  | 35928 |  |  |  | 809.9 | 596.7 |  |  | ~14:00 | 307.1 | 305.6 | 305.5 | 305.7 | 306.8 | 304.5 | 305.5
2023-03-16 |  | sMATF mirror, record drift | 20.441 | 23.892 | 21.7805 | 23.8222 | 24.13 | 23.72 | [5 empty] | 1,76 |  |  |  | 40843 |  |  |  | 809.7 | 597.5 |  | turning RM controller back ON after this measurement | ~14:10 | 307.1 | 305.6 | 305.2 | 305.7 | 308.4 | 304.5 | 305.6
2023-03-16 |  | sMATF mirror, record drift | 20.441 | 23.892 | 21.7805 | 23.8222 | 24.13 | 23.72 | [5 empty] | 1,76 |  |  |  | 41849 |  |  |  | 808.7 | 599.6 |  |  | ~14:20 | 307.2 | 305.6 | 305.4 | 305.8 | 309.2 | 304.6 | 305.7
2023-03-16 |  | sMATF mirror, record drift | 20.441 | 23.892 | 21.7805 | 23.8222 | 24.13 | 23.72 | [5 empty] | 1,76 |  |  |  | 42916 |  |  |  | 808.8 | 602.6 |  |  | ~14:30 | 307.4 | 305.7 | 305.5 | 306 | 309.7 | 304.7 | 305.9
2023-03-16 |  | sMATF mirror, record drift | 20.441 | 23.892 | 21.7805 | 23.8222 | 24.13 | 23.72 | [5 empty] | 1,76 |  |  |  | 43925 |  |  |  | 809 | 603.9 |  | camera off after this measurement to avoid 310K limit, now turning it off after each measurement | ~14:40 | 307.7 | 305.8 | 305.6 | 307.2 | 310 | 304.8 | 306.1
2023-03-16 |  | sMATF mirror, record drift | 20.441 | 23.892 | 21.7805 | 23.8222 | 24.13 | 23.72 | [5 empty] | 1,76 |  |  |  | 44836 |  |  |  | 809.8 | 604.6 |  |  | ~14:50 | 308 | 305.8 | 305.6 | 306.5 | 308.5 | 304.9 | 306.3
2023-03-16 |  | sMATF mirror, record drift | 20.441 | 23.892 | 21.7805 | 23.8222 | 24.13 | 23.72 | [5 empty] | 1,76 |  |  |  | 45806 |  |  |  | 810 | 605.4 |  |  | ~15:00 | 308.3 | 305.9 | 305.7 | 306.6 | 309.5 | 305 | 306.4
2023-03-16 |  | sMATF mirror, record drift | 20.441 | 23.892 | 21.7805 | 23.8222 | 24.13 | 23.72 | [5 empty] | 1,76 |  |  |  | 50835 |  |  |  | 810.6 | 605.6 |  |  | ~15:10 | 308.6 | 306 | 305.7 | 306.8 | 310.3 | 305.1 | 306.6
2023-03-16 |  | sMATF mirror, record drift | 20.441 | 23.892 | 21.7805 | 23.8222 | 24.13 | 23.72 | [5 empty] | 1,76 |  |  |  | 52001 |  |  |  | 810 | 604.7 |  |  | ~15:20 | 308.9 | 306.1 | 305.9 | 306.1 | 308.7 | 305.2 | 306.8
2023-03-16 |  | sMATF mirror, record drift | 20.441 | 23.892 | 21.7805 | 23.8222 | 24.13 | 23.72 | [5 empty] | 1,76 |  |  |  | 52830 |  |  |  | 810.7 | 604.8 |  |  | 15:30 | 309 | 306.2 | 305.9 | 307.1 | 308.3 | 305.3 | 306.8
2023-03-16 |  | sMATF mirror, record drift | 20.441 | 23.892 | 21.7805 | 23.8222 | 24.13 | 23.72 | [5 empty] | 1,76 |  |  |  | 53951 |  |  |  | 810.9 | 605.4 |  |  | 15:40 | 309.2 | 306.2 | 306 | 307.2 | 307.6 | 305.3 | 307
2023-03-16 |  | sMATF retro | 20.441 | 23.892 | 21.7805 | 23.8222 | 24.13 | 23.72 | [7 empty] | 34 | 4.0317 | [7 empty] | RM controller off at 16:00
2023-03-16 |  | sMATF mirror, FARM Off, record drift | 20.441 | 23.892 | 21.7805 | 23.8222 | 24.445 | 23.47 | [5 empty] | 1,76 |  |  |  | 64926 |  |  |  | 810.9 | 595.9 |  |  | 16:50
2023-03-16 |  | sMATF mirror, FARM Off, record drift | 20.441 | 23.892 | 21.7805 | 23.8222 | 24.445 | 23.47 | [5 empty] | 1,76 |  |  |  | 70056 |  |  |  | 810.7 | 595.9 |  |  | 17:00
2023-03-16 |  | sMATF mirror, FARM Off, record drift | 20.441 | 23.892 | 21.7805 | 23.8222 | 24.445 | 23.47 | [5 empty] | 0,19 |  |  |  | 70547 |  | -0.0035 | -0.0013
2023-03-16 |  | sMATF mirror, FARM Off, record drift | 20.441 | 23.892 | 21.7805 | 23.8222 | 24.445 | 23.47 | [5 empty] | 1,76 |  |  |  | 70758 |  |  |  | 810.1 | 595.9 |  |  | 17:09
2023-03-16 |  | sMATF mirror, FARM Off, record drift | 20.441 | 23.892 | 21.7805 | 23.8222 | 24.445 | 23.47 | [5 empty] | 1,77 |  |  |  | 73336 |  |  |  | 810.7 | 595.8 |  |  | 17:35
2023-03-16 |  | PATB mirror | 20.441 | 23.892 | 21.7805 | 23.8222 | 24.445 | 23.47 | 101.501 | 125.706 | -0.004 | 0.1105 | 0.0663 | 0,31 |  |  |  | 75813 |  | -0.0056 | 0.0008
2023-03-16 |  | PATB LED pri | 20.441 | 23.892 | 21.7805 | 23.8222 | 24.445 | 23.47 | 88.669 | 108.404 | -0.004 | 0.1105 | 0.0663 | 1137 |  |  |  | 81057 |  |  |  | 810 | 596
2023-03-16 |  | PATB Retro Pri | 20.441 | 23.892 | 21.7805 | 23.8222 | 24.445 | 23.47 | 66.6936 | 102.3984 | 0 | 0.1105 | 0.0663 | 1137 |  | 19 | 3.4589
2023-03-16 |  | PATB Retro Redundant | 20.441 | 23.892 | 21.7805 | 23.8222 | 24.445 | 23.47 | 92.905 | 144.902 | 0 | 0.1105 | 0.0663 | 1137 |  | 10 | 3.467
2023-03-16 |  | PATB LED Redundant | 20.441 | 23.892 | 21.7805 | 23.8222 | 24.445 | 23.47 | 111.713 | 148.737 | 0 | 0.1105 | 0.0663 | 1137 |  |  |  | 83004 |  |  |  | 810.8 | 595.8
2023-03-16 |  | PATA mirror | 20.441 | 23.892 | 21.7805 | 23.8222 | 24.445 | 23.47 | -76.105 | 125.707 | -0.001 | 0.0877 | 0.059 | 0,23 |  |  |  | 84210 |  | -0.0049 | -6e-05
2023-03-16 |  | PATA LED pri | 20.441 | 23.892 | 21.7805 | 23.8222 | 24.445 | 23.47 | -60.35 | 108.378 | -0.001 | 0.0879 | 0.059 | 1137 |  |  |  | 85053 |  |  |  | 810.7 | 595.98
2023-03-16 |  | PATA Retro Pri | 20.441 | 23.892 | 21.7805 | 23.8222 | 24.445 | 23.47 | -105.494 | 102.701 | 0 | 0.0879 | 0.059 | 1137 |  | 17 | 3.459
2023-03-16 |  | PATA Retro Redundant, failed, maybe becuase of the Leica or the retro | 20.441 | 23.892 | 21.7805 | 23.8222 | 24.445 | 23.47 |  |  |  | 0.0879 | 0.059 | 1137
2023-03-16 |  | PATA LED Redundant | 20.441 | 23.892 | 21.7805 | 23.8222 | 24.445 | 23.47 | -83.424 | 148.621 | -0.001 | 0.0879 | 0.059 | 1137 |  |  |  | 92600 |  |  |  | 810.8 | 595.8
Date |  |  | RM1 |  | RM2 |  | RM3 |  | 5DOFx (mm) | 5DOFy (mm) | 5DOFz (mm) | 5DOF Rx (deg) | 5DOF Ry (deg) | AC Focus | Cam exp(ms) | Lidar Sig | 11 | .fits | .bmp | AC AZ (°) | AC EL (°) | X centroid (pix) | Y centroid (pix)
2023-03-17 20:46 |  | sMATF mirror, FARM Off, record drift | 20.441 | 23.892 | 21.7805 | 23.8222 | 24.445 | 23.47 | [5 empty] | 1,75 |  |  |  | 4605 |  |  |  | 809.6 | 596.6
2023-03-17 20:51 |  | sMATF mirror, FARM Off, record drift | 20.441 | 23.892 | 21.7805 | 23.8222 | 24.445 | 23.47 | [5 empty] | 0,26 |  |  |  | 4926 |  | -0.0015 | -0.0003
2023-03-17 20:52 |  | sMATF retro | 20.441 | 23.892 | 21.7805 | 23.8222 | 24.445 | 23.47 | [7 empty] | 39 | 4.0316
2023-03-17 |  | PATB mirror | 20.441 | 23.892 | 21.7805 | 23.8222 | 24.445 | 23.47 | 101.501 | 125.706 | -0.004 | 0.106 | 0.0697 | 0,26 |  |  |  | 13552 |  | -0.0028 | 0.0008
2023-03-17 22:01 |  | PATB LED pri | 20.441 | 23.892 | 21.7805 | 23.8222 | 24.445 | 23.47 | 88.876 | 108.753 | -0.004 | 0.106 | 0.0698 | 1125 |  |  |  | 15359 |  |  |  | 809 | 596.9 |  |  | 0.207 | 0.349
2023-03-17 22:11 |  | PATB Retro Pri | 20.441 | 23.892 | 21.7805 | 23.8222 | 24.445 | 23.47 | 66.695 | 102.4 | -0.004 | 0.106 | 0.0698 | 1125 |  | 20 | 3.4589
For these measurements, the ADM cooling loops have been on and the RM motor controller and ADM network switch have been on for ~2 hours. Other components (focus motor, camera, leica) are on for as short as possible and not "warmed up" in any way
2023-03-19 | 20:18? | sMATF retro | presumably still at the same positions from 3/17 | [12 empty] | 67 | 4.0297 | [7 empty] | leica settings were left at 20C/ambient pressure
2023-03-19 | 20:19 | sMATF mirror | [11 empty] | 0,26 |  |  |  | 1900 |  | 0.0031 | 0.0028
2023-03-19 | 20:22 | sMATF LED | [11 empty] | 1,76 |  |  |  | 2218 |  |  |  | 809 | 600.8
2023-03-20 | 13:06 | sMATF retro | presumably still at the same positions from 3/17 | [12 empty] | 70 | 4.0288
2023-03-20 | 13:21 | sMATF mirror | [11 empty] | 0,25 |  |  |  | 31432 |  | 0.0006 | -0.0072 |  |  |  | This image was still pretty blurry; I couldn't improve it. Lets watch this
2023-03-20 | 13:57 | sMATF LED | [11 empty] | 1,78 |  |  |  | 35621 |  |  |  | 794.6 | 586
2023-03-21 | 11:24 | sMATF retro | presumably still at the same positions from 3/17 | [12 empty] | 42 | 4.0266
2023-03-21 | 15:47 | Range sMATF mirror | presumably still at the same positions from 3/17 |  |  |  | 24.436 | 23.426 | [7 empty] | 11 | 4.016
2023-03-21 | 15:59 | Range sMATF mirror |  |  |  |  | 24.436 | 23.431 | [7 empty] | 18 | 4.016 | [7 empty] | The sweet spot is ~10um dia. of RM3
2023-03-21 | 16:03 | Range sMATF mirror |  |  |  |  | 24.436 | 23.436 | [7 empty] | 8 | 4.016
2023-03-21 | 16:10 | Range sMATF mirror |  |  |  |  | 24.431 | 23.431 | [7 empty] | 9 | 4.016
2023-03-21 | 16:14 | Range sMATF mirror |  |  |  |  | 24.441 | 23.431 | [7 empty] | 18 | 4.0159
nearly to cold plateau 1
2023-03-22 | NA | sMATF retro | presumably still at the same positions from 3/17 |  |  |  | 24.445 | 23.47 | [7 empty] | NA | NA | [7 empty] | couldn't quite find a return from the retro, not necessary right now... will find it better later. It seems like this position is actually hitting the flat mirror, not the retro
2023-03-22 | 10:40 | sMATF mirror |  |  |  |  | 24.445 | 23.47 | [5 empty] | 0,28 |  |  |  | 103826 |  | 0.0235 | -0.0142
2023-03-22 | 10:43 | sMATF mirror |  |  |  |  | 24.445 | 23.47 | [5 empty] | 0,28 |  |  |  | 104159 |  | 0.0225 | -0.0114
2023-03-22 | 10:36 | sMATF LED |  |  |  |  | 24.445 | 23.47 | [5 empty] | 1,72 |  |  |  | 103439 |  |  |  | 783.8 | 604.6
2023-03-23 | 0:59 | sMATF mirror | presumably still at the same positions from 3/17 |  |  |  | 24.426 | 23.426 | [7 empty] | 15 | 4.0169
2023-03-23 | 0:32 | sMATF mirror |  |  |  |  | 24.441 | 23.431 | [5 empty] | 0,19 |  |  |  | 2541 |  | 0.0197 | -0.0149
2023-03-23 | 0:22 | sMATF LED |  |  |  |  | 24.441 | 23.431 | [5 empty] | 1,77 |  |  |  | 1828 |  |  |  | 783.6 | 596.8
2023-03-23 | 1:00 | sMATF mirror |  |  |  |  | 24.434 | 23.426 | [7 empty] | 41 | 4.017
2023-03-23 | 2:16 | sMATF mirror |  |  |  |  | 24.424 | 23.443 | [5 empty] | 0,25 |  |  |  | 21600 |  | 0.018 | -0.0142
2023-03-23 | 2:22 | sMATF LED |  |  |  |  | 24.424 | 23.443 | [5 empty] | 1,80 |  |  |  | 21749 |  |  |  | 784.5 | 595.8
2023-03-23 | 3:19 | sMATF mirror |  |  |  |  | 24.424 | 23.443 | [5 empty] | 0,25 |  |  |  | 31626 |  | 0.0187 | -0.0135
2023-03-23 | 3:17 | sMATF LED |  |  |  |  | 24.424 | 23.443 | [5 empty] | 1,80 |  |  |  | 31526 |  |  |  | 784.6 | 594.8
2023-03-23 | 4:57 | sMATF mirror |  |  |  |  | 24.424 | 23.443 | [5 empty] | 0,26 |  |  |  | 45509 |  | 0.0187 | -0.0131
2023-03-23 | 4:59 | sMATF LED |  |  |  |  | 24.424 | 23.443 | [5 empty] | 1,80 |  |  |  | 45738 |  |  |  | 784.6 | 594.8
2023-03-23 | 6:35 | sMATF mirror |  |  |  |  | 24.424 | 23.443 | [5 empty] | 0,26 |  |  |  | 63236 |  | 0.018 | -0.0131
2023-03-23 | 6:32 | sMATF LED |  |  |  |  | 24.424 | 23.443 | [5 empty] | 1,80 |  |  |  | 63047 |  |  |  | 784.6 | 594.8
2023-03-23 | 6:39 | sMATF mirror |  |  |  |  | 24.424 | 23.423 | [7 empty] | 17 | 4.0171
2023-03-23 | 6:43 | sMATF mirror |  |  |  |  | 24.424 | 23.423 | [7 empty] | 25 | 4.017
2023-03-23 | 6:46 | sMATF mirror |  |  |  |  | 24.424 | 23.423 | [7 empty] | 6 | 4.017
Date |  |  | RM1 |  | RM2 |  | RM3 |  | 5DOFx (mm) | 5DOFy (mm) | 5DOFz (mm) | 5DOF Rx (deg) | 5DOF Ry (deg) | AC Focus | Cam exp(ms) | Lidar Sig | Range (m) | .fits | .bmp | AC AZ (°) | AC EL (°) | X centroid (pix) | Y centroid (pix)
2023-03-23 | 9:01 | Range of sMATF mirror |  |  |  |  | 24.424 | 23.423 | [7 empty] | 16 | 4.017
2023-03-23 | 9:35 | sMATF mirror |  |  |  |  | 24.424 | 23.423 | [5 empty] | 0,26 |  |  |  | 93539 |  | 0.0183 | -0.0135 |  |  |  | very foggy image, training only
2023-03-23 | 9:54 | sMATF LED |  |  |  |  | 24.424 | 23.423 | [5 empty] | 1,74 |  |  |  | 95220 |  |  |  | 781.7 | 593.8
2023-03-23 | 10:41 | Range sMATF mirror, trial |  |  |  |  | 24.426 | 23.431 | [7 empty] | 7 | 4.017
2023-03-23 | 10:44 | Range sMATF mirror, optimal |  |  |  |  | 24.431 | 23.431 | [7 empty] | 17 | 4.017 | [7 empty] | Range standard deviation ~3um
2023-03-23 | 15:22 | Range sMATF mirror, optimal |  |  |  |  | 24.425 | 23.425 | [7 empty] | 10 | 4.017 | [7 empty] | Range standard deviation ~5um
2023-03-23 | 15:32 | sMATF retro |  |  |  |  | 24.12 | 23.695 | [7 empty] | 35 | 4.0275
2023-03-23 | 16:05 | sMATF LED |  |  |  |  | 24.424 | 23.423 | [5 empty] | 1,75 |  |  |  | 60615 |  |  |  | 782.6 | 593.67
2023-03-23 | 16:09 | sMATF mirror |  |  |  |  | 24.424 | 23.423 | [5 empty] | 0,25 |  |  |  | 62036 |  | 0.0162 | -0.0142
2023-03-23 | 17:40 | PATB mirror | [6 empty] | 101.506 | 125.713 | 0 | 0.1006 | 0.00894 |  |  |  |  | 73820 |  | 0.0169 | -0.0131
2023-03-23 | 17:48 | Range of PATB mirror |  |  |  |  | 24.422 | 23.423 | 101.506 | 125.713 | 0 | 0.1006 | 0.00894 |  |  | 7 | 3.4686 | [7 empty] | adjust RM3A slighltly
2023-03-23 | 18:00 | PATB LED pri |  |  |  |  | 24.422 | 23.423 | 88.238 | 109.109 | 0 | 0.1008 | 0.0895 | 1138 |  |  |  | 75755 |  |  |  | 782.8 | 594.3
2023-03-23 | 18:04 | PATB Retro Pri |  |  |  |  | 24.12 | 23.695 | 66.69 | 102.399 | 0.001 | 0.1008 | 0.0895 |  |  | 33 | 3.4559
Date |  |  | RM1 |  | RM2 |  | RM3 |  | 5DOFx (mm) | 5DOFy (mm) | 5DOFz (mm) | 5DOF Rx (deg) | 5DOF Ry (deg) | AC Focus | Cam exp(ms) | Lidar Sig | Range (m) | .fits | .bmp | AC AZ (°) | AC EL (°) | X centroid (pix) | Y centroid (pix) |  | 0.54943 | 0.54839
2023-03-24 | 6:20 | sMATF retro | presumably still at the same positions from 3/17 |  |  |  | 24.12 | 23.695 | [7 empty] | 42 | 4.0275 | [8 empty] | -0.001037
2023-03-24 | 6:24 | sMATF mirror |  |  |  |  | 24.12 | 23.695 | [5 empty] | 0,26 |  |  |  | 62151 |  | 0.0159 | -0.0149
2023-03-24 | 6:34 | sMATF redundant LED |  |  |  |  | 24.12 | 23.695 | [5 empty] | 1,79 |  |  |  | 63208 |  |  |  | 782.7 | 593.5 |  | 0.5716
2023-03-24 | 6:39 | sMATF mirror |  |  |  |  | 24.424 | 23.423 | [7 empty] | 23 | 4.017
2023-03-24 | 7:08 | PATB mirror | presumably still at the same positions from 3/17 |  |  |  | 24.424 | 23.423 | 101.508 | 125.713 | 0.001 | 0.1005 | 0.0892 | 0,31 |  | 27 | 3.4686 | 70530 |  | 0.0156 | -0.0117
2023-03-24 | 7:23 | PATB LED pri |  |  |  |  | 24.424 | 23.423 | 88.283 | 109.138 | 0.001 | 0.1006 | 0.0892 | 1133 |  |  |  | 72214 |  |  |  | 782.8 | 593.8
2023-03-24 | 7:31 | PATB LED red |  |  |  |  | 24.424 | 23.423 | 111.24 | 149.378 | 0.001 | 0.1003 | 0.0891 | 1134 | [7 empty] | 783.6 | 593.9
2023-03-24 | 7:41 | PATB retro pri |  |  |  |  | 24.12 | 23.695 | 66.693 | 102.398 | 0.001 | 0.1007 | 0.0892 |  |  | 31 | 3.4559
2023-03-24 | 7:48 | PATB retro red |  |  |  |  | 24.12 | 23.695 | 92.91 | 144.902 | 0.001 | 0.1003 | 0.0891 |  |  | 12 | 3.464
2023-03-24 | 8:08 | PATA mirror | presumably still at the same positions from 3/17 |  |  |  | 24.424 | 23.423 | 76.103 | 125.704 | 0.001 | 0.0713 | 0.0615 | 0,25 |  | 27 | 3.4686 | 80413 |  | 0.0162 | -0.0166
2023-03-24 | 8:54 | PATA LED pri |  |  |  |  | 24.424 | 23.423 | -60.915 | 109.07 | -0.001 | 0.0714 | 0.0614 | 1136 |  |  |  | 85247 |  |  |  | 782.95 | 594.5
2023-03-24 | 9:08 | PATA LED red |  |  |  |  | 24.424 | 23.423 | -83.896 | 149.289 | 0.001 | 0.071 | 0.0614 | 1130 |  |  |  | 90700 |  |  |  | 782.6 | 593.8
2023-03-24 | 9:32 | PATA retro pri |  |  |  |  | 24.12 | 23.695 | -105.498 | 102.699 | -0.001 | 0.0714 | 0.0613 |  |  | 23 | 3.456
2023-03-24 | 9:33 | PATA retro red has no Lidar return |  |  |  |  | 24.12 | 23.695 | -26.99 | 144.601 | 0 | 0.0711 | 0.0614
Date |  |  | RM1 |  | RM2 |  | RM3 |  | 5DOFx (mm) | 5DOFy (mm) | 5DOFz (mm) | 5DOF Rx (deg) | 5DOF Ry (deg) | AC Focus | Cam exp(ms) | Lidar Sig | Range (m) | .fits | .bmp | AC AZ (°) | AC EL (°) | X centroid (pix) | Y centroid (pix)
2023-03-31 | 8:46 | sMATF retro | presumably still at the same positions from 3/17 |  |  |  | Same as 3/24 | [8 empty] | 21 | 4.0274
2023-03-31 |  | sMATF redundant LED | 20.0655 | 243.074 | 22.0191 | 23.4807 | [11 empty] | 14118
for this test, the RM controller and ADM network switch have been on since ~04:45AM
2023-04-01 | 6:50 | sMATF retro | presumably still at the same positions from 3/31 |  |  |  | 24.133 | 23.715 | [7 empty] | 32 | 4.0272 | [7 empty] | leica set to 20C, 0mbar. I wasn't sure what temperature to use :(
2023-04-01 | 6:54 | sMATF mirror AC |  |  |  |  | 24.133 | 23.715 | [5 empty] | 0,29 |  |  |  | 65249 |  | 0.0041 | 0.006
2023-04-01 | 6:57 | sMATF mirror AC |  |  |  |  | 24.133 | 23.715 | [5 empty] | 0,29 |  |  |  | 65557 |  | 0.0055 | 0.006 |  |  |  | making two measurements at different parts of the image, the AC is very blurred but seems repeatable?
2023-04-01 | 7:01 | sMATF primary LED |  |  |  |  | 24.133 | 23.715 | [5 empty] | 1,81 |  |  |  | 65933 |  |  |  | 812.8 | 606.4 |  | primary LED is working! We're close-ish at current RM1/2 positions (would have to home them and start over to do better)
2023-04-01 | 7:21 | sMATF mirror AC |  |  |  |  | 24.133 | 23.715 | [5 empty] | 0,25 |  |  |  | 71941 |  | 0.0041 | 0.006
2023-04-01 | 7:18 | sMATF primary LED |  |  |  |  | 24.133 | 23.715 | [5 empty] | 1,81 |  |  |  | 71658 |  |  |  | 812.7 | 605.5 |  | getting centroid+AC for repeatability
2023-04-01 | 7:24 | sMATF primary LED |  |  |  |  | 24.133 | 23.715 | [5 empty] | 1,77 |  |  |  | 72323 |  |  |  | 812.7 | 604.7 |  | seems like Y is slowly drifting down?
2023-04-01 | 7:41 | sMATF mirror range |  |  |  |  | ? | ? | 26.216 | 42.502 | [14 empty] | can't seem to get a good result after hunting around for ~20min or so...
2023-04-01 | 7:24 | sMATF primary LED |  |  |  |  | 24.133 | 23.715 | [5 empty] | 1,77 |  |  |  | 74302 |  |  |  | 811.9 | 604.6 |  | still drifting a little?
2023-04-01 | 7:49 | sMATF mirror AC |  |  |  |  | 24.133 | 23.715 | [5 empty] | 0,25 |  |  |  | 74901 |  | 0.0038 | 0.0056 |  |  |  | yep Az/El is pretty consistent even with the super fuzzy image
2023-04-01 | 11:25 | PATB mirror |  |  |  |  | 24.133 | 23.715 | 101.507 | 125.713 | 0 | 0.1017 | 0.0859 | 0,28 |  |  |  | 12358 |  | 0.0041 | 0.0035
2023-04-01 | 11:34 | PATB LED pri |  |  |  |  | 24.133 | 23.715 | 88.435 | 109.266 | 0 | 0.1018 | 0.0856 | 1133 |  |  |  | 13429 |  |  |  | 812.6 | 605.9
2023-04-01 | 11:39 | PATB mirror range |  |  |  |  | 24.133 | 23.715 | [16 empty] | Couldn't get any signal; moving onto other targets amid time concerns
2023-04-01 | 11:51 | PATB LED red |  |  |  |  | 24.133 | 23.715 | 111.387 | 149.525 | 0 | 0.1014 | 0.0859 | 1134 |  |  |  | 14641 |  |  |  | 812.8 | 606.7
2023-04-01 | 11:43 | PATB retro pri |  |  |  |  | 24.133 | 23.715 | 66.69 | 102.398 | 0 | 0.1018 | 0.086 |  |  | 31 | 3.4557
2023-04-01 | 11:54 | PATB retro red |  |  |  |  | 24.133 | 23.715 | 92.906 | 144.9 | 0 | 0.1014 | 0.086 |  |  | 27 | 3.4638
2023-04-01 | 12:04 | PATA mirror |  |  |  |  | 24.133 | 23.715 | -76.107 | 125.701 | -0.001 | 0.0772 | 0.0653 | 0,25 |  |  |  | 20130 |  | 0.0037 | 0.0084
2023-04-01 | 12:10 | PATA LED pri |  |  |  |  | 24.133 | 23.715 | -60.679 | 109.222 | -0.002 | 0.0774 | 0.065 | 1131 |  |  |  | 20844 |  |  |  | 812.6 | 606.6
2023-04-01 | 12:16 | PATA LED red |  |  |  |  | 24.133 | 23.715 | -83.718 | 149.404 | -0.001 | 0.0771 | 0.0652 | 1130 |  |  |  | 21450 |  |  |  | 812.6 | 606.6
2023-04-01 | 12:20 | PATA retro pri |  |  |  |  | 24.133 | 23.715 | -105.499 | 102.7 | 0 | 0.0775 | 0.0652 |  |  | 43 | 3.4557
2023-04-01 | 12:28 | PATA retro red |  |  |  |  | 24.133 | 23.715 | -26.991 | 144.601 | 0 | 0.0772 | 0.0652 | [11 empty] | Couldn't get any signal; no time to investigate unfortunately.
These measurements performed at warm plateau after CP#2
2023-04-03 | 8:54 | sMATF retro | looks like we weren't sure where RM1/2 were for this measurement |  |  |  | 24.153 | 23.755 | [7 empty] | 34 | 4.0316 | [7 empty] | 20C, 0mbar
2023-04-03 | 9:54 | sMATF mirror AC | 20.223 | 23.8925 | 21.7712 | 23.5708 | 24.153 | 23.755 | [5 empty] | 0,19 |  |  |  | 95237 |  | 0.002 | 0.0006 |  |  |  | very clear bits (so must be cold distortion on SALSA that causes fuzzy bits when at CP?)
2023-04-03 | 9:56 | sMATF primary LED | 20.223 | 23.8925 | 21.7712 | 23.5708 | 24.133 | 23.715 | [5 empty] | 1,75 |  |  |  | 95531 |  |  |  | 809.7 | 599.7 |  | just a little drift/settling since the measurement we made when nulling on the RM Null tab
2023-04-03 | 10:14 | CHIPs retro | 20.223 | 23.8925 | 21.7712 | 23.5708 | 24.63 | 23.18 | [7 empty] | 7 | 4.7625 | [7 empty] | Hey got signal this time!! 7 is about as much as we've ever gotten
2023-04-03 | 10:16 | sMATF retro | 20.223 | 23.8925 | 21.7712 | 23.5708 | 24.133 | 23.715 | [7 empty] | 60 | 4.0316
2023-04-03 | 10:20 | sMATF mirror range | 20.223 | 23.8925 | 21.7712 | 23.5708 | 24.429 | 23.423 | [7 empty] | 22 | 4.0211 | [7 empty] | And got a good return with low stdev for the mirror!
2023-04-03 | 10:35 | sMATF mirror AC | 20.223 | 23.8925 | 21.7712 | 23.5708 | 24.133 | 23.715 | [5 empty] | 0,20 |  |  |  | 103350 |  | -0.0022 | 0.0042 |  |  |  | Az/El still good
2023-04-03 | 10:30 | sMATF primary LED | 20.223 | 23.8925 | 21.7712 | 23.5708 | 24.133 | 23.715 | [5 empty] | 1,75 |  |  |  | 102952 |  |  |  | 811.7 | 597.6 |  | X,Y centroid still drifting in +x,-y direction
2023-04-03 | 11:13 | PATB mirror AC | 20.223 | 23.8925 | 21.7712 | 23.5708 | 24.133 | 23.715 | 101.508 | 125.714 | 0.003 | 0.1288 | 0.0787 | 0,20 |  |  |  | 111159 |  | -0.0008 | 0.0001 |  |  |  | NOTE: had to switch back to ambient LES pitch config file!!!!
2023-04-03 | 11:30 | PATB mirror range | [6 empty] | 101.508 | 125.714 | 0.003 | ? | ? | [11 empty] | didn't find any return after hunting around in rx/ry a little
2023-04-03 | 11:35 | PATB LED pri | [6 empty] | 88.228 | 108.364 | 0 | 0.1287 | 0.0787 | 1134 |  |  |  | 113510 |  |  |  | 807.2 | 602 |  | LED at nominal
2023-04-03 | 11:39 | PATB LED pri | [6 empty] | 88.526 | 108.364 | 0 | 0.1287 | 0.0788 | 1134 |  |  |  | 113931 |  |  |  | 812.1 | 598.8 |  | LED at close to last sMATF LED position
2023-04-03 | 11:43 | PATB retro pri | [6 empty] | 66.887 | 102.399 | -0.001 | 0.1288 | 0.0788 |  |  | 35 | 3.4585 | [7 empty] | 20C, 0mbar ~5stdev
2023-04-03 | 11:52 | PATB retro red | [6 empty] | 92.91 | 144.902 | 0 | 0.1283 | 0.0789 |  |  | 25 | 3.4666 | [7 empty] | 20C, 0mbar ~5stdev
2023-04-03 | 11:59 | PATB LED red | [6 empty] | 111.535 | 148.726 | 0 | 0.1283 | 0.0788 | 1134 |  |  |  | 115805 |  |  |  | 812.5 | 597.9 |  | LED at nominal
2023-04-03 | 12:04 | PATB LED red | [6 empty] | 111.316 | 148.726 | 0 | 0.1283 | 0.0789 | 1134 |  |  |  | 120156 |  |  |  | 808.1 | 600.9 |  | LED at close to last sMATF LED position
2023-04-03 | 12:13 | PATB mirror range |  |  |  |  | 24.429 | 23.423 | 101.505 | 125.714 | 0 | 0.1285 | 0.0788 |  |  | 49 | 3.4712 | [7 empty] | note that we had to go back to RM3 values from sMATF mirror range, ~10um stde
2023-04-03 | 12:16 | PATB mirror range |  |  |  |  | 24.429 | 23.423 | 101.505 | 125.714 | 0 | 0.1285 | 0.0788 |  |  | 23 | 3.4712 | [7 empty] | didn't actually change anything but did another measurement sent and got better stdev (~5um)
ADM measurements with Clint, Pete, Pat, and Derek on the way to cold plateau 3
2023-04-06 | 9:45 | sMATF mirror AC | 20.1561 | 24.0712 | 22.0288 | 23.5519 | 24.429 | 23.423 | [5 empty] | 0,25 |  |  |  | 94415 |  | 0.0041 | 0.0004 |  |  |  | 094524 and 095625 were AC images with different exposure times (250ms vs 2000ms)
2023-04-06 | 9:57 | sMATF primary LED | 20.1561 | 24.0712 | 22.0288 | 23.5519 | 24.429 | 23.423 | [5 empty] | 1,76 |  |  |  | 95625 |  |  |  | 810.75 | 598.6
2023-04-06 | 10:23 | CHIPs retro | 20.1561 | 24.0712 | 22.0288 | 23.5519 | 24.63 | 23.18 | [7 empty] | 7 | 4.7582
2023-04-06 | 10:20 | sMATF retro | 20.1561 | 24.0712 | 22.0288 | 23.5519 | 24.153 | 23.755 | [7 empty] | 16 | 4.0273
2023-04-06 | 10:09 | sMATF mirror range | 20.1561 | 24.0712 | 22.0288 | 23.5519 | ? | ? | forgot to record the RM3 values for putting measurement beam at 808,602 | [6 empty] | 28 | 4.0168 | [7 empty] | 100143 image is infinity focus with leica measure beam on (and primary LED because why not) 100649 is sMATF focus with leica measurement beam sMATF_mirror_range_initial.txt (~10 um stdev) sMATF_mirror_range_final_2.txt (~3 um stdev) for repeatability
2023-04-06 | 10:29 | sMATF mirror range | 20.1561 | 24.0712 | 22.0288 | 23.5519 | 24.435 | 23.426 | this RM3 puts the measurement beam at roughly 808,602 and gets us good return | [6 empty] | 18 | 4.0168 | [7 empty] | sMATF_mirror_range_final_3.txt (~5 um stdev)
2023-04-06 | 10:41 | sMATF mirror AC | 20.1561 | 24.0712 | 22.0288 | 23.5519 | 24.435 | 23.426 | [5 empty] | 0,20 |  |  |  | 104154 |  | 0.002 | 0.0011 |  |  |  | taking repeated AC and centroid measurements to hopefully see drift settle
2023-04-06 | 10:46 | sMATF primary LED | 20.1561 | 24.0712 | 22.0288 | 23.5519 | 24.435 | 23.426 | [5 empty] | 1,76 |  |  |  | 104305 |  |  |  | 810.8 | 597.8
2023-04-06 | 12:03 | sMATF mirror AC | 20.1561 | 24.0712 | 22.0288 | 23.5519 | 24.435 | 23.426 | [5 empty] | 0,20 |  |  |  | 120335 |  | 0.0027 | 0.0039
2023-04-06 | 1200 | sMATF primary LED | 20.1561 | 24.0712 | 22.0288 | 23.5519 | 24.435 | 23.426 | [5 empty] | 1,76 |  |  |  | 120017 |  |  |  | 812.615 | 597.568
2023-04-06 | 13:04 | sMATF mirror AC | 20.1561 | 24.0712 | 22.0288 | 23.5519 | 24.435 | 23.426 | [5 empty] | 0,25 |  |  |  | 130303 |  | 0.0048 | 0.0042
2023-04-06 | 13:08 | sMATF primary LED | 20.1561 | 24.0712 | 22.0288 | 23.5519 | 24.435 | 23.426 | [5 empty] | 1,78 |  |  |  | 130646 |  |  |  | 811.87 | 597.628
2023-04-06 | 14:10 | sMATF mirror AC | 20.1561 | 24.0712 | 22.0288 | 23.5519 | 24.435 | 23.426 | [5 empty] | 0,25 |  |  |  | 140926 |  | 0.0017 | 0.0028
2023-04-06 | 13:58 | sMATF primary LED | 20.1561 | 24.0712 | 22.0288 | 23.5519 | 24.435 | 23.426 | [5 empty] | 1,79 |  |  |  | 135828 |  |  |  | 812.671 | 596.77
2023-04-06 | 15:13 | PATB mirror AC | 20.1561 | 24.0712 | 22.0288 | 23.5519 | 24.435 | 23.426 | 101.51 | 125.714 | 0.001 | 0.0997 | 0.0872 | 0,26 |  |  |  | 151134 |  | 0.0013 | 0.0001
2023-04-06 | 15:11 | PATB mirror range | [6 empty] | 101.51 | 125.714 | 0.001 | 0.0997 | 0.0872 |  |  | 13 | 3.4684
2023-04-06 | 15:45 | PATB LED pri | [6 empty] | 88.51 | 109.275 | 0.002 | 0.0998 | 0.0874 | 1128 |  |  |  | 154505 |  |  |  | 812.743 | 596.808
2023-04-06 | 16:00 | PATB retro pri |  |  |  |  | 24.153 | 24.755 | 66.687 | 102.4 | 0 | 0.0998 | 0.0874 |  |  | 15 | 3.4557
2023-04-06 | 16:30 | PATA mirror AC |  |  |  |  | 24.435 | 23.426 | -76.114 | 125.7 | -0.002 | 0.0848 | 0.0699 |  |  | 19 | 3.4684 | 162826 |  | 0.002 | 0.0022 |  |  |  | Adjusted Rx,Ry to center measurement beam on camera center
2023-04-06 | 16:51 | PATA LED red |  |  |  |  | 24.435 | 23.426 | -83.677 | 149.385 | 0.001 | 0.0771 | 0.0652 | 1128 |  |  |  | 165008 |  |  |  | 812.594 | 597.509
2023-04-06 | 16:55 | PATA retro red |  |  |  |  | 24.153 | 24.755 | did not see return | [6 empty] | 1 | na | [7 empty] | Could not get any laser signal or even any image on camera
2023-04-06 | 17:26 | sMATF mirror | 20.1561 | 24.0712 | 22.0288 | 23.5519 | 24.432 | 23.426 | [5 empty] | 0,20 |  | 18 | 4.0168 | 172617 |  | 0.001 | 0.0018 |  |  |  | Moved RM3A 3um to improve signal
2023-04-06 | 17:33 | sMATF primary LED | 20.1561 | 24.0712 | 22.0288 | 23.5519 | 24.432 | 23.426 | [5 empty] | 1,76 |  |  |  | 173348 |  |  |  | 811.9 | 596.7
2023-04-07 | 9:55 | sMATF mirror | 20.1561 | 24.0712 | 22.0288 | 23.5519 | 24.432 | 23.426 | [5 empty] | 0,20 |  | 20 | 4.0168 | 95551 |  | 0.0041 | 0.0008 |  |  |  | Moved RM3A 3um to improve signal
2023-04-07 | 9:35 | sMATF primary LED | 20.1561 | 24.0712 | 22.0288 | 23.5519 | 24.432 | 23.426 | [5 empty] | 1,73 |  |  |  | 95302 |  |  |  | 810.911 | 599.608
2023-04-07 |  | sMATF retro | 20.1561 | 24.0712 | 22.0288 | 23.5519 | 24.153 | 23.755 | [7 empty] | 18 | 4.0273 | [7 empty] | Standard deviation = 18um
2023-04-07 |  | sMATF retro | 20.1561 | 24.0712 | 22.0288 | 23.5519 | 24.156 | 23.757 | [7 empty] | 16 | 4.0273 | [7 empty] | Moved mirror 2um x,y no signal change, stdev better 5um
2023-04-07 | 10:42 | PATB mirror AC | 20.1561 | 24.0712 | 22.0288 | 23.5519 | 24.432 | 23.426 | 101.51 | 125.714 | 0.001 | 0.0997 | 0.0872 | 0,26 |  |  |  | 104030 |  | 0.0027 | -0.0034 |  |  |  | Used Rx,Ry from this for other PATB measurements
2023-04-07 | 10:40 | PATB mirror range | [6 empty] | 101.51 | 125.714 | 0.001 | 0.1089 | 0.0868 |  |  | 23 | 3.4684 | 110137 |  | 0.0107 | 0.0018 |  |  |  | Adjusted Rx,Ry to peak range signal. Tried to use Rx,Ry from mirror range, but resulted in out of tolerance autocol
2023-04-07 | 11:22 | PATB LED pri | [6 empty] | 88.374 | 109.275 | 0.002 | 0.0992 | 0.0867 | 1128 |  |  |  | 112047 |  |  |  | 810.86 | 599.8
2023-04-07 | 11:31 | PATB retro pri |  |  |  |  | 24.156 | 23.757 | 66.687 | 102.4 | 0 | 0.0992 | 0.0868 |  |  | 7 | 3.4557
2023-04-08 | 11:51 | sMATF mirror | 20.1561 | 24.0712 | 22.0288 | 23.5519 | 24.432 | 23.426 | [5 empty] | 0,24 |  | 15 | 4.0168 | 115001 |  | 0.0034 | 0.0015 |  |  |  | Rehomed RM3
2023-04-08 | 11:58 | sMATF primary LED | 20.1561 | 24.0712 | 22.0288 | 23.5519 | 24.432 | 23.426 | [5 empty] | 1,75 |  |  |  | 115650 |  |  |  | 811.605 | 598.72
2023-04-08 | 12:06 | sMATF retro | 20.1561 | 24.0712 | 22.0288 | 23.5519 | 24.153 | 23.755 | [7 empty] | 15 | 4.0274
2023-04-08 | 12:33 | PATB mirror AC | 20.1561 | 24.0712 | 22.0288 | 23.5519 | 24.432 | 23.426 | 101.51 | 125.715 | 0.002 | 0.0994 | 0.0864 | 0,23 |  |  |  | 121337 |  | 0.0027 | -0.0037 |  |  |  | (no Rx Ry adj)
2023-04-08 | 12:39 | PATB mirror range | [6 empty] | 101.51 | 125.714 | 0.001 | 0.1089 | 0.0868 |  |  | 23 | 3.4684 | 123449
2023-04-08 | 12:47 | PATB LED pri | [6 empty] | 88.374 | 109.275 | 0.002 | 0.0992 | 0.0867 | 1128 |  |  |  | 124633 |  |  |  | 811.5 | 599.763 |  | (no x,y adj)
2023-04-08 | 12:57 | PATB retro pri |  |  |  |  | 24.156 | 23.757 | 66.689 | 102.4 | 0 | 0.0995 | 0.0864 |  |  | 12 | 3.4557
[7 empty] | 24.432 | 23.426 | [5 empty] | 0,0 | [10 empty] | End of test positions
Post CP3 at warm measurements
2023-04-10 | 9:10 | sMATF retro | 20.2274 | 23.8756 | 21.7523 | 23.5699 | 24.153 | 23.755 | [7 empty] | 13 | 4.0317
2023-04-10 | 9:12 | sMATF mirror range | 20.2274 | 23.8756 | 21.7523 | 23.5699 | 24.432 | 23.426 | [7 empty] | 10 | 4.0212
2023-04-10 | 9:13 | sMATF mirror AC | 20.2274 | 23.8756 | 21.7523 | 23.5699 | 24.432 | 23.426 | [5 empty] | 0,23 |  |  |  | 91233 |  | 0.0017 | 0.0011
2023-04-10 | 9:14 | sMATF primary LED | 20.2274 | 23.8756 | 21.7523 | 23.5699 | 24.432 | 23.426 | [5 empty] | 1,75 |  |  |  | 91356 |  |  |  | 809.7 | 598.7
2023-04-10 | 9:25 | sMATF primary LED | 20.2274 | 23.8756 | 21.7523 | 23.5699 | 24.432 | 23.426 | [5 empty] | 1,75 |  |  |  | 92529 |  |  |  | 810.7 | 597.8
2023-04-10 | 10:00 | PATB mirror | 20.2274 | 23.8756 | 21.7523 | 23.5699 | 24.432 | 23.426 | 101.512 | 125.718 | 0.002 | 0.128 | 0.0822 | 0,23 |  | 18 | 3.4712 | 95833 |  | -0.0004 | 0.0014
2023-04-10 | 10:13 | PATB LED pri | 20.2274 | 23.8756 | 21.7523 | 23.5699 |  |  | 88.547 | 108.421 | -0.001 | 0.1282 | 0.0822 | 1128 |  |  |  | 101304 |  |  |  | 810.3 | 597.62
2023-04-10 | 10:21 | PATB retro pri, rm3 slightly diferent from sMATF | 20.2274 | 23.8756 | 21.7523 | 23.5699 | 24.153 | 23.74 | 66.688 | 102.4 | -0.001 | 0.1283 | 0.0822 |  |  | 10 | 3.4585
2023-04-10 | 10:40 | sMATF primary LED | 20.2274 | 23.8756 | 21.7523 | 23.5699 | 24.153 | 23.74 | [5 empty] | 1,75 |  |  |  | 103920 |  |  |  | 810.694 | 597.564 |  | repeatability on sMATF measurements
2023-04-10 | 10:43 | sMATF mirror AC | [11 empty] | 0,23 |  |  |  | 104249 |  | 0.001 | 0.0025
2023-04-10 | 10:46 | Turn off FARM and the RM power supply to avoid overheating | [22 empty] | Turn off FARM and the RM power supply to avoid overheating
2023-04-10 | 14:22 | sMATF retro | should be the same as this morning, had to power off RM controller due to thermal concerns |  |  |  | 24.153 | 23.755 | [7 empty] | 15 | 4.0317 | [7 empty] | post ambient/vacuum FDPR ADM check
2023-04-10 | 14:24 | sMATF mirror range |  |  |  |  | 24.432 | 23.426 | [7 empty] | 18 | 4.0212
2023-04-10 | 14:26 | sMATF mirror AC |  |  |  |  | 24.432 | 23.426 | [5 empty] | 0,23 |  |  |  | 142504 |  | -0.0056 | 0.0011
2023-04-10 | 14:28 | sMATF primary LED |  |  |  |  | 24.432 | 23.426 | [5 empty] | 1,75 |  |  |  | 142710 |  |  |  | 808.8 | 591.5
2023-04-10 | 14:45 | PATB mirror |  |  |  |  | 24.432 | 23.426 | 101.513 | 125.721 | -0.003 | 0.1279 | 0.0819 | 0,23 |  | 26 | 3.4712 | 144436 |  | -0.0022 | -0.0006
2023-04-10 | 14:54 | PATB LED pri | [6 empty] | 88.617 | 108.613 | -0.002 | 0.128 | 0.0819 | 1128 |  |  |  | 145410 |  |  |  | 808.8 | 591.2
2023-04-10 | 14:54 | PATB retro pri |  |  |  |  | 24.153 | 23.755 | 66.689 | 102.399 | -0.003 | 0.128 | 0.0819 |  |  | 5 | 3.4585
Note on E223: [Threaded comment] Your version of Excel allows you to read this threaded comment; however, any edits to it will get removed if the file is opened in a newer version of Excel. Learn more: https://go.microsoft.com/fwlink/?linkid=870924 Comment: I think this value of 21.66 is suspect. I think it was REALLY ~23.6, but the value for RM1A was input again on accident.
Note on AB308: [Threaded comment] Your version of Excel allows you to read this threaded comment; however, any edits to it will get removed if the file is opened in a newer version of Excel. Learn more: https://go.microsoft.com/fwlink/?linkid=870924 Comment: Please check XY-DoF value for cells I304:J305

## Scratch - Extracted Data
Date |  | RM1 |  | RM2 |  | RM3 |  | 5DOFx | 5DOFy | 5DOFz | Rx | Ry | Range(m) | AC AZ | AC EL | X centroid | Y centroid
2023-01-10 | sMATF flat mirror | 21.65 | 23 | 20.49 | 25.225 | 24.4 | 23.45
2023-01-10 | sMATF retro | 21.65 | 23 | 20.49 | 25.225 | 24.1469 | 23.75 | [5 empty] | 4.0317
2023-01-10 | sMATF flat mirror | 21.76 | 23.44 | 21.16 | 25.34 | 24.4 | 23.45 | [6 empty] | 0.008 | -0.002 | 797.2 | 606.7
2023-01-11 | PATB flat mirror | 21.76 | 23.44 | 21.16 | 25.34 | 24.4 | 23.45 | 101.5 | 125.7 |  | 0.08955 | 0.04604 |  | 0.0072 | 0.0004
2023-01-11 | PATB Redundant LED | 21.76 | 23.44 | 21.16 | 25.34 | 24.4 | 23.45 | 111.505 | 148.801 |  | 0.08955 | 0.04604 |  |  |  | 797.7 | 606.1
2023-01-11 | PATB redundant retro | 21.76 | 23.44 | 21.16 | 25.34 | 24.4 | 23.45 | 124.405 | 142.401 |  | 0.08955 | 0.04604 | 3.4668 |  |  | 797.4 | 606.3
2023-01-11 | PATB primary retro | 21.76 | 23.44 | 21.16 | 25.34 | 24.4 | 23.45 | 100.693 | 100.398 |  | 0.08955 | 0.04604 | 3.4588 |  |  | 797.5 | 605.5
2023-01-11 | PATB primary LED | 21.76 | 23.44 | 21.16 | 25.34 | 24.4 | 23.45 | 88.503 | 108.398 |  | 0.08955 | 0.04604 |  |  |  | 797.8 | 606.7
2023-01-11 | PATA flat mirror | 21.76 | 23.44 | 21.16 | 25.34 | 24.4 | 23.45 | -76.1 | 125.7 |  | 0.06382 | 0.03495 |  | 0.0062 | -0.0044
2023-01-11 | PATA primary LED | 21.76 | 23.44 | 21.16 | 25.34 | 24.4 | 23.45 | -60.597 | 108.399 |  | 0.06382 | 0.03495 |  |  |  | 797.2 | 606.7
2023-01-11 | PATA primary retro | 21.76 | 23.44 | 21.16 | 25.34 | 24.4 | 23.45 | -72.502 | 100.199 |  | 0.06382 | 0.03495 | 3.4588 |  |  | 795.5 | 607.5
2023-01-11 | PATA redundant retro | 21.76 | 23.44 | 21.16 | 25.34 | 24.4 | 23.45 | ? | 142.101? |  | 0.06382 | 0.03495 | 3.4668 |  |  | 796.4 | 606.5
2023-01-11 | PATA redundant LED | 21.76 | 23.44 | 21.16 | 25.34 | 24.4 | 23.45 | -83.618 | 148.642 |  | 0.06382 | 0.03495 |  |  |  | 797 | 606.7
2023-01-11 | PATB flat mirror | 21.76 | 23.44 | 21.16 | 25.34 | 24.4 | 23.45 | 101.5 | 125.7 |  | 0.08766 | 0.04135 |  | 0.0062 | -0.0048
2023-01-21 | PATA Mirror | 21.76 | 23.44 | 21.16 | 25.34 | 24.4 | 23.45 | -76.1 | 125.7 |  | 0.06382 | 0.03495 |  | 0.0041 | -0.002
2023-01-21 | PATA Primary LED | 21.76 | 23.44 | 21.16 | 25.34 | 24.4 | 23.45 | -60.656 | 108.448 |  | 0.06382 | 0.03495 |  |  |  | 796.9 | 607.9
2023-01-21 | PATA Primary Retro | 21.76 | 23.44 | 21.16 | 25.34 | 24.4 | 23.45 | -71.5 | 100.2 |  | 0.06382 | 0.03495 | 3.4588 |  |  | 781 | 620
2023-01-21 | PATA redantant Retro | 21.76 | 23.44 | 21.16 | 25.34 | 24.4 | 23.45 | -97.508 | 142.6 |  | 0.06382 | 0.03495 | 3.4669 |  |  | 775 | 613
2023-01-21 | PATA redundant LED | 21.76 | 23.44 | 21.16 | 25.34 | 24.4 | 23.45 | -83.76 | 148.69 |  | 0.06382 | 0.03495 |  |  |  | 797.2 | 607
2023-01-21 | PATB Mirror | 21.76 | 23.44 | 21.16 | 25.34 | 24.4 | 23.45 | 101.54 | 125.7 |  | 0.08955 | 0.04604 |  | 0.0076 | 0.0011
2023-01-21 | PATB Primary LED | 21.76 | 23.44 | 21.16 | 25.34 | 24.4 | 23.45 | 88.353 | 108.404 |  | 0.08955 | 0.04604
2023-01-21 | PATB Primary Retro | 21.76 | 23.44 | 21.16 | 25.34 | 24.4 | 23.45 | 99.702 | 102.4 |  | 0.08955 | 0.04604 | 3.4592 | 759 | 581
2023-01-21 | PATB redundant Retro | 21.76 | 23.44 | 21.16 | 25.34 | 24.4 | 23.45 | 123.415 | 144.4 |  | 0.08955 | 0.04604 | 3.4672
2023-01-21 | PATB redundant LED | 21.76 | 23.44 | 21.16 | 25.34 | 24.4 | 23.45 | 111.354 | 140.801 |  | 0.08955 | 0.04604 |  |  |  | 796.1 | 606.7
2023-01-21 | sMATF flat mirror | 21.76 | 23.44 | 21.16 | 25.34 | 24.4 | 23.45 | [6 empty] | 0.008 | 0.022 | 808.4 | 605.5
2023-01-21 | sMATF Retro | 21.76 | 23.44 | 21.16 | 25.34 | 24.15 | 23.75 | [5 empty] | 4.0322 |  |  | 131 | 1058
2023-01-21 | PATA redundant LED | 21.76 | 23.44 | 21.16 | 25.34 | 24.15 | 23.75 | -83.37 | 148.34 |  | 0.06382 | 0.03495 |  |  |  | 808.1 | 606.5
2023-01-21 | PATA Primary LED | 21.76 | 23.44 | 21.16 | 25.34 | 24.15 | 23.75 | -60.3 | 108.15 |  | 0.06382 | 0.03495 |  |  |  | 808.9 | 605
2023-01-21 | PATB Primary LED | 21.76 | 23.44 | 21.16 | 25.34 | 24.15 | 23.75 | 88.75 | 108.15 |  | 0.08955 | 0.04604 |  |  |  | 808 | 605.7
2023-01-21 | PATB redundant LED | 21.76 | 23.44 | 21.16 | 25.34 | 24.15 | 23.75 | 111.7 | 148.5 |  | 0.08955 | 0.04604 |  |  |  | 807.7 | 606
2023-01-22 | PATB redundant LED | 21.76 | 23.44 | 21.16 | 25.34 | 24.15 | 23.75 | 111.7 | 148.5 |  | 0.08955 | 0.04604 |  |  |  | 807 | 607.9
2023-01-23 | sMATF mirror | 21.76 | 23.44 | 21.16 | 25.34 | 24.15 | 23.75 | [6 empty] | 0.026 | 0.067 | 840 | 614
2023-01-23 | sMATF mirror | 21.75 | 23.32 | 21.04 | 25.31 | 24.15 | 23.75 | [6 empty] | 0.004 | 0.011 | 785 | 597
2023-01-23 | sMATF Retro | 21.75 | 23.32 | 21.04 | 25.31 | 24.15 | 23.75 | [5 empty] | 4.0322 |  |  | 134 | 1036
2023-01-23 | PATB Mirror | 21.75 | 23.32 | 21.04 | 25.31 | 24.15 | 23.75 | 101.5 | 125.7 |  | 0.08955 | 0.04604 |  | -0.004 | -0.014
2023-01-23 | PATB primary LED | 21.75 | 23.32 | 21.04 | 25.31 | 24.15 | 23.75 | 89.4 | 107.3 |  | 0.08955 | 0.04604 |  |  |  | 806.8 | 606.6
2023-01-24 | sMATF mirror (H2) | 21.75 | 23.32 | 21.04 | 25.31 | 24.15 | 23.75 | [6 empty] | -0.0261 | -0.043 | 738.8 | 572.7
2023-01-25 | sMATF mirror (H2) | 21.75 | 23.32 | 21.04 | 25.31 | 24.15 | 23.75 | [6 empty] | -0.03 | 0.016 | 787 | 571
2023-01-25 | sMATF mirror (H2) | 21.68 | 21.66 | 21.53 | 25.27 | 24.15 | 23.75 | [6 empty] | -0.0004 | -0.005 | 796.8 | 599
2023-01-25 | sMATF retro (H2) | 21.68 | 21.66 | 21.53 | 25.27 | 24.15 | 23.75 | [5 empty] | 4.0321
2023-01-25 | PATB mirror | 21.68 | 21.66 | 21.53 | 25.27 | 24.15 | 23.75 | 101.5 | 125.7 |  | 0.089 | 0.046 |  | -0.007 | -0.03
2023-01-25 | PATB mirror | 21.68 | 21.66 | 21.53 | 25.27 | 24.15 | 23.75 | 101.5 | 125.7 |  | 0.106 | 0.061 |  | -0.0011 | -0.0107
2023-01-25 | PATB primary LED | 21.68 | 21.66 | 21.53 | 25.27 | 24.15 | 23.75 | 88.69 | 108.5 |  | 0.106 | 0.061 |  |  |  | 796.8 | 598.6
2023-01-25 | PATB primary Retro | 21.68 | 21.66 | 21.53 | 25.27 | 24.15 | 23.75 | 69.2 | 102.9 |  | 0.106 | 0.061 | 3.4596
2023-01-25 | sMATF mirror | 21.68 | 21.66 | 21.53 | 25.27 | 24.15 | 23.75 | [6 empty] | 0.0013 | -0.0034
2023-01-25 | sMATF mirror | 21.68 | 21.66 | 21.53 | 25.27 | 24.15 | 23.75 | [8 empty] | 798.7 | 599.7
2023-01-25 | sMATF Retro | 21.68 | 21.66 | 21.53 | 25.27 | 24.15 | 23.75 | [5 empty] | 4.0322 |  |  | 154 | 1044 | 3.2e-05
2023-01-25 | PATB mirror, nominal rx and ry | 21.68 | 21.66 | 21.53 | 25.27 | 24.15 | 23.75 | 101.5 | 125.7 |  | 0.106 | 0.061 |  | -0.0011 | -0.01
2023-01-25 | PATB primary LED | 21.68 | 21.66 | 21.53 | 25.27 | 24.15 | 23.75 | 88.69 | 108.4 |  | 0.106 | 0.061 |  |  |  | 797.7 | 600
2023-01-25 | PATB primary Retro | 21.68 | 21.66 | 21.53 | 25.27 | 24.15 | 23.75 | 69.2 | 102.9 |  | 0.106 | 0.061 | 3.4596 |  |  | 219 | 932 | -2e-06
2023-01-26 | sMATF mirror | 21.68 | 21.66 | 21.53 | 25.27 | 24.15 | 23.75 | [6 empty] | 0.0024 | -0.0041
2023-01-26 | sMATF mirror | 21.68 | 21.66 | 21.53 | 25.27 | 24.15 | 23.75 | [8 empty] | 797.6 | 600.4
2023-01-26 | sMATF Retro | 21.68 | 21.66 | 21.53 | 25.27 | 24.15 | 23.75 | [5 empty] | 4.0322 |  |  | 150.5 | 1046.5
2023-01-26 | PATB mirror, nominal rx and ry | 21.68 | 21.66 | 21.53 | 25.27 | 24.15 | 23.75 | 101.503 | 125.705 |  | 0.106 | 0.061 |  | -0.0001 | -0.0107
2023-01-26 | PATB primary LED | 21.68 | 21.66 | 21.53 | 25.27 | 24.15 | 23.75 | 88.695 | 108.398 |  | 0.106 | 0.061 |  |  |  | 797.45 | 599.7
2023-01-26 | PATB primary Retro | 21.68 | 21.66 | 21.53 | 25.27 | 24.15 | 23.75 | 69.2 | 102.898 |  | 0.106 | 0.061 | 3.4596 |  |  | 220 | 933
2023-01-26 | sMATF mirror (5DOF stowed) | 21.68 | 21.66 | 21.53 | 25.27 | 24.15 | 23.75 | [6 empty] | 0.0027 | -0.0065
2023-01-26 | sMATF mirror (5DOF stowed) | 21.68 | 21.66 | 21.53 | 25.27 | 24.15 | 23.75 | [8 empty] | 795.6 | 600.7
2023-01-26 | sMATF Retro (5DOF stowed) | 21.68 | 21.66 | 21.53 | 25.27 | 24.15 | 23.75 | [5 empty] | 4.0322 |  |  | 150 | 1048
2023-01-30 | sMATF mirror, H2 position | 21.68 | 21.66 | 21.53 | 25.27 | 24.15 | 23.75 | [6 empty] | 0.0037 | -0.0058
2023-01-30 | sMATF mirror, pinhole | 21.68 | 21.66 | 21.53 | 25.27 | 24.15 | 23.75 | [8 empty] | 795.7 | 601
2023-01-30 | sMATF Retro | 21.68 | 21.66 | 21.53 | 25.27 | 24.15 | 23.75 | [5 empty] | 4.0321 |  |  | 150 | 1047
2023-01-30 | PATB mirror, nominal rx and ry | 21.68 | 21.66 | 21.53 | 25.27 | 24.15 | 23.75 | 101.503 | 125.705 |  | 0.106 | 0.061 |  | -0.0001 | -0.0104
2023-01-30 | PATB primary LED | 21.68 | 21.66 | 21.53 | 25.27 | 24.15 | 23.75 | 88.595 | 108.4 |  | 0.106 | 0.061 |  |  |  | 795.7 | 601.1
2023-01-30 | PATB primary Retro | 21.68 | 21.66 | 21.53 | 25.27 | 24.15 | 23.75 | 69.2 | 102.9 |  | 0.106 | 0.061 | 3.4596 |  |  | 220 | 931
2023-01-30 | sMATF mirror (5DOF stowed) | 21.68 | 21.66 | 21.53 | 25.27 | 24.15 | 23.75 | [6 empty] | 0.0024 | -0.0065
2023-01-30 | sMATF mirror (5DOF stowed) | 21.68 | 21.66 | 21.53 | 25.27 | 24.15 | 23.75 | [8 empty] | 795.6 | 600.7
2023-01-30 | sMATF Retro (5DOF stowed) | 21.68 | 21.66 | 21.53 | 25.27 | 24.15 | 23.75 | [5 empty] | 4.0322 |  |  | 148 | 1046
2023-01-31 | sMATF mirror, H2 position | 21.68 | 21.66 | 21.53 | 25.27 | 24.15 | 23.75 | [6 empty] | 0.0017 | -0.0048
2023-01-31 | sMATF mirror, pinhole | 21.68 | 21.66 | 21.53 | 25.27 | 24.15 | 23.75 | [8 empty] | 796 | 599.9
2023-01-31 | sMATF Retro | 21.68 | 21.66 | 21.53 | 25.27 | 24.15 | 23.75 | [5 empty] | 4.0322 |  |  | 148 | 1046
2023-02-01 | PATB mirror | 21.68 | 21.66 | 21.53 | 25.27 | 24.15 | 23.75 | 101.5 | 125.7 |  | 0.105 | 0.061 |  | 0.0006 | 0.012
2023-02-01 | PATB primary LED | 21.68 | 21.66 | 21.53 | 25.27 | 24.15 | 23.75 | 88.7 | 108.5 |  | 0.105 | 0.061 |  |  |  | 795.8 | 599.8
2023-02-01 | PATB primary Retro | 21.68 | 21.66 | 21.53 | 25.27 | 24.15 | 23.75 | 69.2 | 102.9 |  | 0.105 | 0.061 | 3.4595 |  |  | 218 | 934
2023-02-02 | sMATF mirror, H2 position | 21.68 | 21.66 | 21.53 | 25.27 | 24.15 | 23.75 | 250 | -65 | -60.0012 | -0.0003 | 0.0003 |  | 0.0048 | -0.0048 |  |  |  | Average sMATF mirror Az/El since 1/25 (deg): | 0.002325 | -0.0051125 | std. dev (arcsec) | 5.616 | 3.9644
2023-02-02 | sMATF mirror, pinhole, H2 position | 21.68 | 21.66 | 21.53 | 25.27 | 24.15 | 23.75 | 250 | -65 | -60.0012 | -0.0003 | 0.0003 |  |  |  | 795.724 | 601.921 |  | Average sMATF LED X/Y since 1/25 (pixels) | 796.4655 | 600.4151 | std. dev (pix) | 1.1491 | 0.89055
2023-02-02 | sMATF Retro, H2 position | 21.68 | 21.66 | 21.53 | 25.27 | 24.15 | 23.75 | 250 | -65 | -60.0012 | -0.0003 | 0.0003 | 4.0321 |  |  | 151 | 1049 |  | Average sMATF retro range since 1/25 (m) | 4.0322 |  | std. dev (microns) | 39.6836
2023-02-02 | PATB mirror, nominal rx=0.105 and ry=0.061 | 21.68 | 21.66 | 21.53 | 25.27 | 24.15 | 23.75 | 101.501 | 125.703 | 0.0034 | 0.10547 | 0.06164 |  | -0.0001 | -0.011 |  |  |  | Average PATB mirror Az/El since 1/25 (deg): | -0.0005 | -0.01056 | std. dev (arcsec) | 1.9718 | 1.3614 | 2/1 AC EL data removed; it seems to be an outlier. Missing negative sign, perhaps?
2023-02-02 | PATB primary LED | 21.68 | 21.66 | 21.53 | 25.27 | 24.15 | 23.75 | 88.596 | 108.399 | 0.0005 | 0.10556 | 0.0617 |  |  |  | 795.707 | 601.771 |  | Average PATB LED X/Y since 1/25 (pixels) | 796.5262 | 600.1618 | std. dev (pix) | 0.91512 | 1.1207
2023-02-02 | PATB primary Retro | 21.68 | 21.66 | 21.53 | 25.27 | 24.15 | 23.75 | 69.2002 | 102.898 | -0.00098 | 0.10596 | 0.06107 | 3.4596 |  |  | 220 | 931 |  | Average PATB retro range since 1/25 (m) | 3.4596 |  | std. dev (microns) | 29.5138
[19 empty] | Average sMATF Az/El since 1/25 (deg) (H2): | 0.00225 | -0.00465 | std. dev (arcsec) | 6.6097 | 2.9489
[19 empty] | Average sMATF X/Y since 1/25 (pixels) (H2) | 796.754 | 600.3202 | std. dev (pix) | 1.2038 | 1.033
[19 empty] | Average sMATF range since 1/25 (m) (H2) | 4.0322 |  | std. dev (microns) | 40.6136
[19 empty] | Average sMATF Az/El since 1/25 (deg) (Stow): | 0.00255 | -0.0065 | Deltas from H2 mean (arcsec): | 1.08 | -6.66
[19 empty] | Average sMATF X/Y since 1/25 (pixels) (Stow) | 795.6 | 600.7 | Deltas from H2 mean (pixels): | -1.154 | 0.37983
[19 empty] | Average sMATF range since 1/25 (m) (Stow) | 4.0322 |  | Delta from H2 mean (um): | 37.6667

## Scratch - RM worksheet
Date | RM1 |  | RM2 |  | RM3 |  | AC Focus | .fits | AC AZ | AC EL | X centroid | Y centroid
2023-02-08 | 21.71 | 23.42 | 21.13 | 24.5 | 24.12 | 23.77 |  |  | 0.32 | 0.034 | 787 | 625
2023-02-08 | 21.71 | 23.42 | 21.13 | 25 |  |  |  |  | 0.1 | 0.01 | 801 | 612
| 21.91 | 23.42 | 21.13 | 25.2 |  |  | 0,17 |  | 0.15 | 0.01 | 802 | 637
| 21.91 | 23.42 | 21.13 | 25.1 |  |  |  |  | -0.008 | 0.0087 | 800 | 509
| 21.81 | 23.42 | 21.13 | 25 |  |  |  |  | -0.04 | 0.01
| 21.81 | 23.42 | 21.13 | 25.02 |  |  |  |  | -0.02 | 0.01
| 21.81 | 23.42 | 21.13 | 25.03 |  |  |  |  | 0 | 0 |  | 516
| 21.34 | 23.42 | 21.13 | 24.73 | [7 empty] | 675
| 21.38 | 23.42 | 21.13 | 24.73 |  |  |  |  | 0.09 | 0.017 |  | 625
|  | 23.42 | 21.13 | 24.67 |  |  |  |  | 0.02
| 21.38 | 23.42 | 21.13 | 24.66 |  |  |  |  | 0 |  |  | 554
| -0.43 |  |  | -0.37 | 0 |  |  | 0 | 0 |  |  | 38
| 20.95 | 23.42 | 21.13 | 24.29 |  |  |  | 0 | 0.0076 | 0.014 | 804.3 | 590
| 20.75 | 23.42 | 21.13 | 24.29 |  |  |  |  | 0.31 | 0.014 | 805 | 851
| 20.95 | 23.32 | 21.13 | 24.29 |  |  |  |  | 0.027 | -0.15 | 664 | 606
| 20.95 | 23.42 | 20.93 | 24.29 |  |  |  |  | 0.024 | 0.24 | 990.5 | 604
| 20.95 | 23.42 | 21.13 | 24.19 |  |  |  |  | -0.15 | 0.007 | 798 | 764
| 20.95 | 23.42 | 21.15 | 24.28 |  |  |  |  | -0.0063 | -0.018 | 777 | 579 |  | 793 | 604
| 20.93 | 23.44 | 21.16 | 24.28 |  |  |  |  | 0.023 | 0.005 | 798 | 606
| 20.942 | 23.428 | 21.147 | 24.274 |  |  |  |  | -0.005 | -0.0013 | 790.2 | 582.9 | -21.1
| -0.23876 |  |  | -0.20545
| 21.1808 |  |  | 24.4794 |  |  |  |  | -0.007 | 0.0008 | 794 | 563
| 20.7032 | 23.428 | 21.147 | 24.0686 |  |  |  |  | 0.0008 | 0.0004 | 793 | 602
Date | RM1 |  | RM2 |  | RM3 |  | AC Focus | .fits | AC AZ | AC EL | X centroid | Y centroid
2023-02-14 | 20.77 | 23.485 | 21.201 | 24.12 | 24.12 | 23.77 |  |  | 0.022 | 0.032 | 808 | 601.5 | [5 empty] | 0.346 | 0.2215
| 20.77 | 23.485 | 21.201 | 24.06 | 24.12 | 23.77 |  |  | -0.082 | 0.032
| 20.697 | 23.55 | 21.301 | 24.06 | 24.12 | 23.77 |  |  | -0.082 | -0.092 | 808 | 602 |  | 0.187 | 0.0595 | 0.072 | -0.12
| 20.697 | 23.55 | 21.301 | 24.06 | 24.12 | 23.77 |  |  | 0.016 | 0.025 | 808 | 602 |  | 0.087 | 0.1595 | -0.066 | 0.07
| 20.57 | 23.83 | 21.701 | 23.94 | 24.12 | 23.77 |  | 81032 |  |  | 807.7 | 601.6
| 20.57 | 23.83 | 21.701 | 23.94 | 24.12 | 23.77 |  | 81341 | 0.005 | 0.004 |  |  |  | RX | RY | AZ | EL | d AZ | d EL
[14 empty] | 0.087 | 0.0595 | -0.008 | -0.014
[14 empty] | 0.187 | 0.0595 | 0.073 | 0.047 | 0.081 | 0.061
[14 empty] | 0.087 | 0.1595 | -0.064 | -0.263 | -0.056 | -0.249
[14 empty] | 0.097 | 0.0515 | 0.006 | -0.012
[14 empty] | 0.09 | 0.0485 | 0.002 | -0.018
[14 empty] | 0.09 | 0.04 | 0.007 | -0.026
[14 empty] | 0.09 | 0.055 | -0.002 | -0.013
[14 empty] | 0.09 | 0.06 | -0.005 | -0.008

## Scratch - notes
Date | Item | RM1 |  | RM2 | RM3
| 1~2 sec current draw | -5000
diable/enable
| 1~2 sec current draw | -5000
| 1~2 sec current draw | -10000
| 1~2 sec current draw |  | -10000
y
2022-12-18
Steps to get an Autocollimation image measurement if you can't use the 'autocollimate' button
1) save a .fits image by right clicking the image and choosing .fits
2) from a command prompt type:
a) 'convert ' (that's a space after the convert)
b) then drag the newly created .fits image from the windows browser (C:\OSIM_TEST_DATA …) to the command line to auto fill it's path
c) then type ' -rotate 90 ' (those are spaces before the -rotate and after the 90)
d) then drag the same .fits filename from the windows browser into the command line
e) you should see something like: "C:\Documents and Settings\osw_dev>convert C:\OSIM_TEST_DATA\ADM_yyyymmddHHMMSS.fits -rotate 90 C:\OSIM_TEST_DATA\ADM_yyyymmddHHMMSS.fits"
f) hit 'enter' which will rotate the fits image 90 degrees
3) again from the command prompt type:
a) 'convert C:\OSIM_TEST_DATA\ADM_yyyymmddHHMMSS.fits C:\OSIM_TEST_DATA\ADM_yyyymmddHHMMSS.bmp' (you can drag and drop the filename from the browser again, just change it to .bmp in the second one)
b) hit 'enter' which will create a .bmp from the rotated .fits image
4) copy the newly created .bmp and paste it into C:\OSIM\abmc\autocollimate
5) open the file 'atac.ini' found in C:\OSIM\abmc\autocollimate
6) replace the currently listed .bmp with the new one you just created (should be the same path, just change the .bmp file name)
7) Open the folder 'C:\OSIM\bin' and double click 'OSIM ADM AT Image Processor.exe' (or if it's already running that's OK too)
8) hit the big green "NEW ATAC IMAGE" button at the top left (when you do it should now list your created .bmp just to the right of the button and it should also now be shown in the image display
9) Using the left/right and up/down errors near the bottom left of the program, position the yellow cross hairs in the zoomed image such that they line up with the horizontal and vertical lines somewhere in the image where the code bits are nice and bright
10) Using the mouse, left click in the zoomed image in the vertical row of columns that have MATCHING code bit values (so the code bit pattern should be repeating in that vertical column)
11) Toward the bottom right of the program, click the little code bit check mark buttons until the red markers match up with the bright code bits in BOTH cells
12) Hit the 'decode' button and you'll have your Az and El measurement
Some directionality notes:
RM | Focus | result
1A | 2118 | increase to move code bit image UP
1B | 2118 | increase to move code bit image RIGHT
2A | 2118 | increase to move code bit image LEFT
2B | 2118 | increase to move code bit image DOWN
2A | 0,24 | increase to move autocollimation elevation NEGATIVE
Target Options (do not delete)
sMATF mirror
sMATF LED pri
sMATF LED red
sMATF retro
CHIPS retro
Aperture mirror
PATB mirror
PATB LED pri
PATB LED red
PATB retro pri
PATB retro red
PATA mirror
PATA LED pri
PATA LED red
PATA retro pri
PATA retro red
MATF mirror
MATF LED pri
MATF LED red
MATF retro
Other (specify in notes)

## ADM keybinds
ADM software shortcuts
F2 | Camera On/Off
F3 | Grab aka "Live"
F4 | Save FITS
F5 | Set focus to Exposure on video tab (now sets to video)
F6 | Crosshair overlay
F7 | Spot measure
F8 | ACQuire Image (F2,F3,F4,F2)
F9 | AC LED1 On/Off
F10 | AC LED2 On/Off
F11 | AutoCol measure
F12 | LDM Power On/Off
F8,F7 | To Acquire Image of Pinole or LED, and then measure Spot Centroid
F9,F10,F8,F9,F10,F11 | To Acquire bit image and then measure auto collimation